{"text_format":"3","workbook":"xlsx","n_sheets":3,"sheets":["Detailed Technical Spec","Sheet2","Sheet3"],"not_rendered":{"images":1,"defined_names":1},"read":"whole","date_system":"1900"}

{"sheet":"Detailed Technical Spec","index":0,"n_rows":540,"n_cols":3,"rows_kept":540,"merged":[[0,1,0,2],[2,0,24,0],[25,0,47,0],[48,0,70,0],[71,0,93,0],[94,0,108,0],[109,0,124,0],[125,0,140,0],[141,0,156,0],[157,0,161,0],[162,0,166,0],[167,0,170,0],[171,0,174,0],[175,0,178,0],[179,0,183,0],[184,0,187,0],[188,0,191,0],[192,0,210,0],[211,0,230,0],[231,0,250,0],[251,0,271,0],[272,0,292,0],[293,0,313,0],[314,0,334,0],[335,0,358,0],[359,0,382,0],[383,0,406,0],[407,0,430,0],[431,0,447,0],[448,0,452,0],[453,0,471,0],[472,0,479,0],[480,0,494,0],[495,0,511,0],[512,0,527,0],[528,0,531,0],[532,0,535,0],[536,0,539,0]]}
[null,"ANNEXURE A: DETAILED TECHNICAL SPECIFICATION REQUIREMENTS\nBID RT40-2024: SUPPLY AND DELIVERY OF CRUTCHES AND WALKING AIDS TO THE STATE FOR THE PERIOD OF 36 MONTHS"]
["Item Number","Item Description","Comply/ Do Not Comply"]
["RT40-01-001","Crutches, elbow (figure 1) of CKS 112"]
[null,"For paediatric use "]
[null,"Extruded anodised aluminium"]
[null,"Upper tube size:  ±25.4 x ±1.6mm"]
[null,"Lower tube size:  ±21.6 x ±1.4mm"]
[null,"The supplier should specify what size ferrule fit this crutch "]
[null,"Injection-moulded polypropylene of grade 2340PC hand grip"]
[null,"Immovable ergonom+C54:C62ically designed hand grip firmly connected to the main frame"]
[null,"18 adjusting holes on the main tube for handle adjustable in 25mm increments which can suit various users."]
[null,"Maximum height between hand grip and floor:  ±710mm (C)"]
[null,"Height between cuff and hand grip:  ±120 mm (A)"]
[null,"Minimum height of hand grip from end: ±340mm (B)"]
[null,"Weight:  ±600g"]
[null,"Injection-moulded poly propylene of grade 2340PC cuff, "]
[null,"hinged to allow a 90 degree movement in the vertical plane with no movement in the horizontal plane"]
[null,"Removable non marking non-slip ferrule with steel insert"]
[null,"Injection-moulded rubber PVC ferrule"]
[null,"Double pin spring clip type of at least 6mm in diameter electroplated to prevent corrosion"]
[null,"Adjust by two spring-loaded pins (spring clips) on each foot piece"]
[null,"Foot piece to slide over the full extensibility and can easily disassembled clearance between sling parts not to exceed ±1mm"]
[null,"Able to carry a weight up to 150kg by a full pair use"]
[null,"Must comply to the latest issue of CKS 112"]
[null,"Pack size pair"]
["RT40-01-002","Crutches, elbow (figure 1) of CKS 112"]
[null,"For child use "]
[null,"Extruded anodised aluminium"]
[null,"Upper tube size:  ±25.4 x ±1.6mm"]
[null,"Lower tube size:  ±21.6 x ±1.4mm"]
[null,"Injection-moulded polypropylene of grade 2340PC hand grip"]
[null,"The supplier should specify what size ferrule fit this crutch "]
[null,"Immovable ergonomically designed hand grip firmly connected to the main frame"]
[null,"18 adjusting holes on the main tube for handle adjustable in 25mm increments which can suit various users."]
[null,"Maximum height between hand grip and floor:  ±780mm (C)"]
[null,"Height between cuff and hand grip:  ±200 mm (A)"]
[null,"Minimum height of hand grip from end: ±380mm (B)"]
[null,"Weight:  ±600g"]
[null,"Injection-moulded poly propylene of grade 2340PC cuff, "]
[null,"hinged to allow a 90 degree movement in the vertical plane with no movement in the horizontal plane"]
[null,"Removable non marking non-slip ferrule with steel insert"]
[null,"Injection-moulded rubber PVC ferrule"]
[null,"Double pin spring clip type of at least 6mm in diameter electroplated to prevent corrosion"]
[null,"Adjust by two spring-loaded pins (spring clips) on each foot piece"]
[null,"Foot piece to slide over the full extensibility and can easily disassembled clearance between sling parts not to exceed ±1mm"]
[null,"Able to carry a weight up to 150kg by a full pair use"]
[null,"Must comply to the latest issue of CKS 112"]
[null,"Pack size pair"]
["RT40-01-003","Crutches, elbow (figure 1) of CKS 112"]
[null," For adult use "]
[null,"Extruded anodised aluminium"]
[null,"Upper tube size:  ±25.4 x ±1.6mm"]
[null,"Lower tube size:  ±21.6 x ±1.4mm"]
[null,"Injection-moulded polypropylene of grade 2340PC hand grip"]
[null,"The supplier should specify what size ferrule fit this crutch "]
[null," Immovable ergonomically designed hand grip firmly connected to the main frame"]
[null,"18 adjusting holes on the main tube for handle adjustable in 25mm increments which can suit various users."]
[null," Maximum height between hand grip and floor:  ±940mm (C)"]
[null,"Height between cuff and hand grip:  ±250 mm (A)"]
[null,"Minimum height of hand grip from end: ±530mm (B)"]
[null,"Weight:  ±600g"]
[null,"Injection-moulded poly propylene of grade 2340PC cuff, "]
[null,"hinged to allow a 90 degree movement in the vertical plane with no movement in the horizontal plane"]
[null,"Removable non marking non-slip ferrule with steel insert"]
[null," Injection-moulded rubber PVC ferrule"]
[null,"Double pin spring clip type of at least 6mm in diameter electroplated to prevent corrosion"]
[null,"Adjust by two spring-loaded pins (spring clips) on each foot piece"]
[null,"Foot piece to slide over the full extensibility and can easily disassembled clearance between sling parts not to exceed ±1mm"]
[null,"Able to carry a weight up to 150kg by a full pair use"]
[null,"Must comply to the latest issue of CKS 112"]
[null,"Pack size pair"]
["RT40-01-004","Crutches, elbow, Heavy Duty (Figure 1) of CKS 112"]
[null,"For adult use "]
[null,"Stainless steel"]
[null,"Upper tube size:  ±25.4 x ±1.6mm"]
[null,"Lower tube size:  ±21.6 x ±1.4mm"]
[null,"The supplier should specify what size ferrule fit this crutch "]
[null,"Injection-moulded polypropylene of grade 2340PC hand grip"]
[null,"Immovable ergonomically designed hand grip firmly connected to the main frame"]
[null,"18 adjusting holes on the main tube for handle adjustable in 25mm increments which can suit various users."]
[null,"Maximum height between hand grip and floor:  ±940mm (C)"]
[null,"Height between cuff and hand grip:  ±250 mm (A)"]
[null,"Minimum height of hand grip from end: ±530mm (B)"]
[null,"Weight:  ±600g"]
[null,"Injection-moulded poly propylene of grade 2340PC cuff, "]
[null,"hinged to allow a 90 degree movement in the vertical plane with no movement in the horizontal plane"]
[null,"Removable non marking non-slip ferrule with steel insert"]
[null,"Injection-moulded rubber PVC ferrule"]
[null,"Double pin spring clip type of at least 6mm in diameter electroplated to prevent corrosion"]
[null,"Adjust by two spring-loaded pins (spring clips) on each foot piece"]
[null,"Foot piece to slide over the full extensibility and can easily disassembled clearance between sling parts not to exceed ±1mm"]
[null,"Able to carry a weight up to ±120kg per crutch"]
[null,"Must comply to the latest issue of CKS 112"]
[null,"Pack size pair"]
["RT40--02-001","Crutches, Axilla (Figure 2)  OF SANS 491 TYPE 2"]
[null,"Laminated saligna wood "]
[null,"Size 1: Paediatric (kiddies small)"]
[null,"Adjustable high:  725 (D) to 825mm (E)"]
[null,"Height top part between hand grip and end:  305mm (C)"]
[null,"Height of lower part:   350mm (A)"]
[null,"Handgrips:  injection-moulded  polypropylene grade 2340PC adjustable hand grip, non-slip handles without finger grooves with one treaded Metal pins with butterfly screw and washer"]
[null,"Non marking non-slip injection moulded rubber ferrule with steel insert"]
[null,"Injection moulded rubber PVC ferrule"]
[null,"Adjustable preformed bow with rectangular cross-section with two treaded Metal pins with butterfly screw and washer"]
[null,"18 adjusting holes on the main tube for handle adjustable in 25mm increments which can suit various users."]
[null,"Axilla  pads: soft grey foam-like"]
[null,"Length of axilla pad:  ±160mm (B)"]
[null,"Must comply to the latest issue of SANS 491 TYPE 2"]
[null,"Pack size pair "]
["RT40-02-002","Crutches, Axilla (Figure 2)  OF SANS 491 TYPE 2"]
[null,"Laminated saligna wood "]
[null,"Size 2: Child"]
[null,"Adjustable high:  850 (D) to 1000mm (E)"]
[null,"height top part between hand grip and end:  325mm (C)"]
[null,"height of lower part:   400mm (A)"]
[null,"Handgrips:  injection-moulded grey polypropylene grade 2340PC adjustable hand grip, non-slip handles without finger grooves with one treaded Metal pins with butterfly screw and washer"]
[null,"non marking non-slip ferrule with steel insert"]
[null," Injection moulded rubber PVC ferrule"]
[null,"adjustable preformed bow with rectangular cross-section with two treaded Metal pins with butterfly screw and washer"]
[null,"18 adjusting holes on the main tube for handle adjustable in 25mm increments which can suit various users."]
[null,"Axilla  pads: soft grey foam-like"]
[null,"Length of axilla pad:  ±160mm (B)"]
[null,"The supplier should specify what size ferrule fit this crutch "]
[null,"Must comply to the latest issue of SANS 491 TYPE 2"]
[null,"Pack size pair"]
["RT40-02-003","Crutches, Axilla (Figure 2)  OF SANS 491 TYPE 2"]
[null,"Laminated saligna wood "]
[null,"Size 3: Adult ( Medium)"]
[null,"Adjustable high:  1025 (D) to 1250mm (E)"]
[null,"Height top part between hand grip and end:  340mm (C)"]
[null,"Height of lower part:   470mm (A)"]
[null,"Handgrips:  injection-moulded polypropylene grade 2340PC adjustable hand grip, non-slip handles without finger grooves with one treaded Metal pins with butterfly screw and washer"]
[null,"Non marking non-slip ferrule with steel insert"]
[null," PVC Injection moulded rubber ferrule "]
[null,"Adjustable preformed bow with rectangular cross-section with two treaded Metal pins with butterfly screw and washer"]
[null,"18 adjusting holes on the main tube for handle adjustable in 25mm increments which can suit various users."]
[null," Axilla  pads: soft foam-like"]
[null,"Length of axilla pad:  ±205mm (B)"]
[null,"The supplier should specify what size ferrule fit this crutch "]
[null,"Must comply to the latest issue of SANS 491 TYPE 2"]
[null,"Pack size pair"]
["RT40-02-004","Crutches, Axilla (Figure 2)  OF SANS 491 TYPE 2"]
[null,"Laminated saligna wood "]
[null,"Size 4: Large Adult"]
[null,"Adjustable high:  1275 (D) to 1500mm (E)"]
[null,"Height top part between hand grip and end:  430mm (C)"]
[null,"Height of lower part:   470mm (A)"]
[null,"Handgrips:  injection-moulded polypropylene grade 2340PC adjustable hand grip, non-slip handles without finger grooves with one treaded Metal pins with butterfly screw and washer"]
[null,"Non marking non-slip ferrule with steel insert"]
[null,"Injection moulded rubber PVC ferrule"]
[null,"Adjustable preformed bow with rectangular cross-section with two treaded Metal pins with butterfly screw and washer"]
[null,"18 adjusting holes on the main tube for handle adjustable in 25mm increments which can suit various users."]
[null,"Axilla  pads: soft foam-like"]
[null,"Length of axilla pad:  ±205mm (B)"]
[null,"The supplier should specify what size ferrule fit this crutch "]
[null,"Must comply to the latest issue of SANS 491 TYPE 2"]
[null,"Pack size pair"]
["RT40-02-005","Arm pad to be used with axilla crutch"]
[null,"soft foam-like "]
[null,"Must fit a Size 1 - Peadiatric and Size 2 - Child Axilla Crutch."]
[null,"Must be compatible with the items awarded for size 1 and 2 axilla crutch"]
[null,"Pack size:  each"]
["RT40-02-006","Arm pad to be used with axilla crutch"]
[null,"Soft foam-like."]
[null,"Must fit a Size 3 - Adult (medium) and Size 4 - Adult (large) Axilla Crutch."]
[null,"Must be compatible with the items awarded for size 3 and 4 axilla crutch"]
[null,"Pack size:  each"]
["RT40-03-001","Tips/ferrules with washer "]
[null,"Moulded plasticised PVC, non-marking, non-slip with steel insert, removable"]
[null,"Size:  ±16mm inner diameter "]
[null,"Each"]
["RT40-03-002","Tips/ferrules with washer "]
[null,"Moulded rubber PVC, non-marking, non-slip with steel insert, removable"]
[null,"Size:  ±18mm inner diameter "]
[null,"Each "]
["RT40-03-003","Tips/ferrules with washer "]
[null,"Moulded rubber PVC, non-marking, non-slip with steel insert, removable"]
[null,"Size:  ±20mm inner diameter "]
[null,"each"]
["RT40-03-004","Tips/ferrules with washer "]
[null,"Moulded plasticised PVC, non-marking, non-slip with steel insert, removable"]
[null,"Size:  ±22mm inner diameter "]
[null,"The tip/ferrules must fit and be compatible with Crutches, elbow (adult use)"]
[null,"Each "]
["RT40-03-005","Tips/ferrules with washer "]
[null,"Moulded rubber PVC, non-marking, non-slip with steel insert, removable"]
[null,"Size:  ±25mm inner diameter "]
[null,"Each"]
["RT40-03-006","Tips/ferrules with washer "]
[null,"Moulded rubber PVC, non-marking, non-slip with steel insert, removable"]
[null,"This item must fit and be compatible with the items awarded in size 1, 2, 3, and 4 axilla crutch"]
[null,"Pack size:  each"]
["RT40-04-001","Crutches, Gutter (Figure 3)  of CKS 112"]
[null,"Child use"]
[null,"Extruded anodised aluminium crutch, "]
[null,"Upper tube size:  ±25.4 x ±1.6mm"]
[null,"Lower tube size:  ±21.6 x ±1.4mm "]
[null,"The supplier should specify what size ferrule fit this crutch "]
[null,"Height adjustable between: 720-940mm (B),"]
[null,"Height top part between hand grip and end ±535mm (A), "]
[null,"18 adjusting holes on the main tube for handle adjustable in 25mm increments which can suit various users. "]
[null,"Weight ±1kg"]
[null,"Cuff/Gutter:  padded upholstery cover, adjustable by means of velcro straps. "]
[null,"Adjustability between hand grip and gutter to be between :  ±255 – 380mm (C) "]
[null,"Handgrip: Injection- moulded polypropylene of grade 2340PC, immovable ergonomically designed, firmly connected to the main frame, "]
[null,"Ferrule/tip: Removable non-marking, non-slip ferrule with steel insert, injection-moulded  rubber PVC "]
[null,"Double pin spring clip type at least 6mm in diameter electroplated to prevent corrosion, "]
[null,"Adjust by two spring- loaded pins (spring clips) on each foot piece, "]
[null,"Foot pieces to slide freely over the full extensibility and can be easily disassembled clearance between sliding parts not to exceed 1mm"]
[null,"Must comply to the latest issue of CKS 112"]
[null,"Each"]
["RT40-04-002","Crutches, Gutter (Figure 3)  of CKS 112"]
[null,"Adult use"]
[null,"Extruded anodised aluminium crutch, "]
[null,"Upper tube size:  ±25.4 x ±1.6mm"]
[null,"Lower tube size:  ±21.6 x ±1.4mm"]
[null,"The supplier should specify what size ferrule fit this crutch "]
[null,"Injection- moulded polypropylene of grade 2340PC hand grip, "]
[null,"Height adjustable between: ±915-1140mm (B),"]
[null,"Height top part between hand grip and end ±705mm (A), "]
[null,"18 adjusting holes on the main tube for handle adjustable in 25mm increments which can suit various users. "]
[null,"Weight ±1kg"]
[null,"Cuff/Gutter: padded upholstery cover, adjustable by means of velcro straps to secure forearm and elbow in place"]
[null,"Adjustability between hand grip and gutter to be between 320 and  gutter between:  ±320 – 450mm C"]
[null,"Handgrip: injection- moulded polypropylene of grade 2340PC, immovable ergonomically designed, firmly connected to the main frame, "]
[null,"Ferrule/tip: Removable non-marking, non-slip ferrule with steel insert, injection-moulded rubber PVC "]
[null,"Double pin spring clip type at least 6mm in diameter electroplated to prevent corrosion, "]
[null,"Adjust by two spring- loaded pins (spring clips) on each foot piece, "]
[null,"Foot pieces to slide freely over the full extensibility and can be easily disassembled clearance between sliding parts not to exceed 1mm"]
[null,"Must comply to the latest issue of CKS 112"]
[null,"Each"]
["RT40-04-003","Rollator, gutter,"]
[null,"For adult use,"]
[null,"extruded natural anodised aluminium non-folding frame"]
[null,"Tubing ±21.6x ±1.4mm, 6063T6 aluminium tubing, "]
[null,"Handgrip: injection-moulded rubber PVC immovable ergonomically designed, "]
[null,"Height adjustable: 960-1015mm, "]
[null,"Width: 385-400mm, "]
[null,"Distance between back legs: 500-505mm, "]
[null,"Length of gutter arm support: 255-265mm, "]
[null,"Castor: ±75mm non marking, front"]
[null,"The supplier should specify what size ferrule fit this item"]
[null,"Maximum weight: 3kg, "]
[null,"Adjusted height by pressing single or double spring-loaded pins(springs clips)on each foot piece, 6mm diameter, electroplated to prevent corrosion, "]
[null,"Foot pieces to slide freely over the full extensibility and can be easily disassembled, 1mm minimum clearance between sliding parts, "]
[null,"Distance between middle of gutters: ±385mm, "]
[null,"Ferrule/tip:  removable injection-moulded , rubber PVC non-slip, non-marking fitted with steel insert, "]
[null,"Gutter length adjustable, "]
[null,"Gutter: Padded upholstery cover, tension over forearm adjustable by means of a hook and loop fasteners"]
[null,"Must comply to the latest issue of CKS 112 for aluminium part"]
[null,"Each"]
["RT40-05-001","Pulpit Walker, Standard, Folding, without wheels"]
[null,"For paediatric use"]
[null,"Extruded anodised aluminium"]
[null,"Handgrip: Immovable ergonomically designed, injection-moulded rubber P.V.C "]
[null,"Adjustable Height: 350 – 500mm, "]
[null,"Width between handles: ±300mm"]
[null,"The supplier should specify what size ferrule fit this item"]
[null,"Maximum weight: ±2kg"]
[null,"Cross bars between front and rear legs"]
[null,"Folding mechanism:  Angled Cross-bar between front legs with swivel mechanism to fold frame. Straight cross-bar at top with depression of single spring-loaded pin (spring clip) to fold device."]
[null,"8 - 10 adjusting holes on the legs for handle adjustable in 25mm increments which can suit various users. "]
[null,"Depression of single or double spring-loaded pins (spring clips) on each foot piece to adjust, "]
[null,"Foot pieces to slide freely over the full extensibility and can be easily disassembled, "]
[null,"1mm minimum clearance between sliding parts, "]
[null,"Ferrule/tip removable non-slip, non-marking fitted with steel insert, injection-moulded rubber PVC, "]
[null,"Standard, double or single pin type spring clip at least 6mm diameter, electroplated to prevent corrosion"]
[null,"Folded Depth:  ±110mm"]
[null,"The supplier should specify what size ferrule fit this item"]
[null,"Must comply to the latest issue of CKS 112 for aluminium part"]
[null,"Size: Extra Small"]
[null,"Each"]
["RT40-05-002","Pulpit Walker, Standard, Folding, without wheels"]
[null,"For child use"]
[null,"Extruded anodised aluminium"]
[null,"Handgrip: Immovable ergonomically designed, injection-moulded rubber P.V.C "]
[null,"Adjustable Height: 450 – 600mm, "]
[null,"Width between handles: ±350mm"]
[null,"The supplier should specify what size ferrule fit this item"]
[null,"Maximum weight: ±2kg"]
[null,"Cross bars between front and rear legs"]
[null,"Folding mechanism:  Angled Cross-bar between front legs with swivel mechanism to fold frame. Straight cross-bar at top with depression of single spring-loaded pin (spring clip) to fold device."]
[null,"8 - 10 adjusting holes on the legs for handle adjustable in 25mm increments which can suit various users. "]
[null,"Depression of single or double spring-loaded pins (spring clips) on each foot piece to adjust, "]
[null,"Foot pieces to slide freely over the full extensibility and can be easily disassembled, "]
[null,"1mm minimum clearance between sliding parts, "]
[null,"Ferrule/tip removable non-slip, non-marking fitted with steel insert, injection-moulded rubber PVC, "]
[null,"Standard, double or single pin type spring clip at least 6mm diameter, electroplated to prevent corrosion"]
[null,"Folded Depth:  ±110mm"]
[null,"The supplier should specify what size ferrule fit this item"]
[null,"Must comply to the latest issue of CKS 112 for aluminium part"]
[null,"Size: Small"]
[null,"Each"]
["RT40-05-003","Pulpit Walker, Standard, Folding, without wheels"]
[null,"For adult/child use"]
[null,"Extruded anodised aluminium"]
[null,"Handgrip: Immovable ergonomically designed, injection-moulded rubber P.V.C "]
[null,"Adjustable Height: 550 – 750mm, "]
[null,"Width between handles: ±400mm"]
[null,"The supplier should specify what size ferrule fit this crutch "]
[null,"Maximum weight: ±2kg"]
[null,"Cross bars between front and rear legs"]
[null,"Folding mechanism:  Angled Cross-bar between front legs with swivel mechanism to fold frame. Straight cross-bar at top with depression of single spring-loaded pin (spring clip) to fold device."]
[null,"8 - 10 adjusting holes on the legs for handle adjustable in 25mm increments which can suit various users. "]
[null,"Depression of single or double spring-loaded pins (spring clips) on each foot piece to adjust, "]
[null,"Foot pieces to slide freely over the full extensibility and can be easily disassembled, "]
[null,"1mm minimum clearance between sliding parts, "]
[null,"Ferrule/tip removable non-slip, non-marking fitted with steel insert, injection-moulded rubber PVC, "]
[null,"Standard, double or single pin type spring clip at least 6mm diameter, electroplated to prevent corrosion"]
[null,"Folded Depth:  ±110mm"]
[null,"The supplier should specify what size ferrule fit this crutch "]
[null,"Must comply to the latest issue of CKS 112 for aluminium part"]
[null,"Size: Medium"]
[null,"Each"]
["RT40-05-004","Pulpit Walker, Standard, Folding, without wheels"]
[null,"For adult use"]
[null,"Extruded anodised aluminium"]
[null,"Handgrip: Immovable ergonomically designed, injection-moulded rubber P.V.C "]
[null,"Adjustable Height: 700 – 900mm, "]
[null,"Width between handles: ±500mm"]
[null,"The supplier should specify what size ferrule fit this crutch "]
[null,"Maximum weight: ±2kg"]
[null,"Cross bars between front and rear legs"]
[null,"Folding mechanism:  Angled Cross-bar between front legs with swivel mechanism to fold frame. Straight cross-bar at top with depression of single spring-loaded pin (spring clip) to fold device."]
[null,"8 - 10 adjusting holes on the legs for handle adjustable in 25mm increments which can suit various users. "]
[null,"Depression of single or double spring-loaded pins (spring clips) on each foot piece to adjust, "]
[null,"Foot pieces to slide freely over the full extensibility and can be easily disassembled, "]
[null,"1mm minimum clearance between sliding parts, "]
[null,"Ferrule/tip removable non-slip, non-marking fitted with steel insert, injection-moulded rubber PVC, "]
[null,"Standard, double or single pin type spring clip at least 6mm diameter, electroplated to prevent corrosion"]
[null,"Folded Depth:  ±110mm"]
[null,"The supplier should specify what size ferrule fit this crutch "]
[null,"Must comply to the latest issue of CKS 112 for aluminium part"]
[null,"Size: Large"]
[null,"Each"]
["RT40-05-005","Pulpit Walker, Folding, with wheels/castors"]
[null,"Size 1: For Paediatric use"]
[null,"Extruded anodised aluminium"]
[null,"Handgrip: Immovable ergonomically designed, injection-moulded rubber P.V.C "]
[null,"Adjustable Height: 350 – 650mm, "]
[null,"Width between handles: ±300mm"]
[null,"The supplier should specify what size ferrule fit this item"]
[null,"Maximum weight: ±2kg"]
[null,"Cross bars between front and rear legs"]
[null,"Folding mechanism:  Angled Cross-bar between front legs with swivel mechanism to fold frame. Straight cross-bar at top with depression of single spring-loaded pin (spring clip) to fold device."]
[null,"8 - 10 adjusting holes on the legs for handle adjustable in 25mm increments which can suit various users. "]
[null,"Depression of single or double spring-loaded pins (spring clips) on each foot piece to adjust, "]
[null,"Foot pieces to slide freely over the full extensibility and can be easily disassembled, "]
[null,"1mm minimum clearance between sliding parts, "]
[null,"Ferrule/tip removable non-slip, non-marking fitted with steel insert, injection-moulded rubber PVC, "]
[null,"Standard, double or single pin type spring clip at least 6mm diameter, electroplated to prevent corrosion"]
[null,"Folded Depth:  ±110mm"]
[null,"The supplier should specify what size ferrule fit this item"]
[null,"Wheels/castors connected to front legs of frame"]
[null,"Wheels/castor size:  ±75mm in diameter, non-marking, rubberized"]
[null,"The supplier should specify what size ferrule fit this item"]
[null,"Must comply to the latest issue of CKS 112 for aluminium part"]
[null,"Size: Extra Small"]
[null,"Each"]
["RT40-05-006","Pulpit Walker, Folding, with wheels/castors"]
[null,"For child use"]
[null,"Extruded anodised aluminium"]
[null,"Handgrip: Immovable ergonomically designed, injection-moulded rubber P.V.C "]
[null,"Adjustable Height: ±450 – 600mm, "]
[null,"Width between handles: ±350mm"]
[null,"The supplier should specify what size ferrule fit this item"]
[null,"Maximum weight: ±2kg"]
[null,"Cross bars between front and rear legs"]
[null,"Folding mechanism:  Angled Cross-bar between front legs with swivel mechanism to fold frame. Straight cross-bar at top with depression of single spring-loaded pin (spring clip) to fold device."]
[null,"8 - 10 adjusting holes on the legs for handle adjustable in 25mm increments which can suit various users. "]
[null,"Depression of single or double spring-loaded pins (spring clips) on each foot piece to adjust, "]
[null,"Foot pieces to slide freely over the full extensibility and can be easily disassembled, "]
[null,"1mm minimum clearance between sliding parts, "]
[null,"Ferrule/tip removable non-slip, non-marking fitted with steel insert, injection-moulded rubber PVC, "]
[null,"Standard, double or single pin type spring clip at least 6mm diameter, electroplated to prevent corrosion"]
[null,"Folded Depth:  ±110mm"]
[null,"The supplier should specify what size ferrule fit this item"]
[null,"Wheels/castors connected to front legs of frame"]
[null,"Wheels/castor size:  ±75mm in diameter, non-marking, rubberized"]
[null,"The supplier should specify what size ferrule fit this item"]
[null,"Must comply to the latest issue of CKS 112 for aluminium part"]
[null,"Size: Small"]
[null,"Each"]
["RT40-05-007","Pulpit Walker, Folding, with wheels/castors"]
[null,"For adult/child use"]
[null,"Extruded anodised aluminium"]
[null,"Handgrip: Immovable ergonomically designed, injection-moulded rubber P.V.C "]
[null,"Adjustable Height: 550 – 750mm, "]
[null,"Width between handles: ±400mm"]
[null,"The supplier should specify what size ferrule fit this item"]
[null,"Maximum weight: ±2kg"]
[null,"Cross bars between front and rear legs"]
[null,"Folding mechanism:  Angled Cross-bar between front legs with swivel mechanism to fold frame. Straight cross-bar at top with depression of single spring-loaded pin (spring clip) to fold device."]
[null,"8 - 10 adjusting holes on the legs for handle adjustable in 25mm increments which can suit various users. "]
[null,"Depression of single or double spring-loaded pins (spring clips) on each foot piece to adjust, "]
[null,"Foot pieces to slide freely over the full extensibility and can be easily disassembled, "]
[null,"1mm minimum clearance between sliding parts, "]
[null,"Ferrule/tip removable non-slip, non-marking fitted with steel insert, injection-moulded rubber PVC, "]
[null,"Standard, double or single pin type spring clip at least 6mm diameter, electroplated to prevent corrosion"]
[null,"Folded Depth:  ±110mm"]
[null,"The supplier should specify what size ferrule fit this item"]
[null,"Wheels/castors connected to front legs of frame"]
[null,"Wheels/castor size:  ±75mm in diameter, non-marking, rubberized"]
[null,"The supplier should specify what size ferrule fit this item"]
[null,"Must comply to the latest issue of CKS 112 for aluminium part"]
[null,"Size: Medium"]
[null,"Each"]
["42211506-00008","Pulpit Walker, Folding, with wheels/castors"]
[null,"For adult use"]
[null,"Extruded anodised aluminium"]
[null,"Handgrip: Immovable ergonomically designed, injection-moulded rubber P.V.C "]
[null,"Adjustable Height: 700 – 900mm, "]
[null,"Width between handles: ±500mm"]
[null,"The supplier should specify what size ferrule fit this item"]
[null,"Maximum weight: ±2kg"]
[null,"Cross bars between front and rear legs"]
[null,"Folding mechanism:  Angled Cross-bar between front legs with swivel mechanism to fold frame. Straight cross-bar at top with depression of single spring-loaded pin (spring clip) to fold device."]
[null,"8 - 10 adjusting holes on the legs for handle adjustable in 25mm increments which can suit various users. "]
[null,"Depression of single or double spring-loaded pins (spring clips) on each foot piece to adjust, "]
[null,"Foot pieces to slide freely over the full extensibility and can be easily disassembled, "]
[null,"1mm minimum clearance between sliding parts, "]
[null,"Ferrule/tip removable non-slip, non-marking fitted with steel insert, injection-moulded rubber PVC, "]
[null,"Standard, double or single pin type spring clip at least 6mm diameter, electroplated to prevent corrosion"]
[null,"Folded Depth:  ±110mm"]
[null,"The supplier should specify what size ferrule fit this item"]
[null,"Wheels/castors connected to front legs of frame"]
[null,"Wheels/castor size:  ±75mm in diameter, non-marking, rubberized"]
[null,"The supplier should specify what size ferrule fit this item"]
[null,"Must comply to the latest issue of CKS 112 for aluminium part"]
[null,"Size: Large"]
[null,"Each"]
["RT40-05-009","Reverse walker, "]
[null,"For paediatric/child use"]
[null,"Folding, natural anodised aluminium frame, "]
[null,"Handgrips: injection-moulded rubber P.V.C, immovable ergonomically designed"]
[null,"Height adjustable:  450 - 525mm"]
[null,"Width: ±490mm width"]
[null,"The supplier should specify what size ferrule fit this item"]
[null,"Distance between back legs: ±85mm, "]
[null,"Distance between front castors: ±537mm, "]
[null,"Castors: ±75mm non marking castors, "]
[null,"The supplier should specify what size ferrule fit this item"]
[null,"Depression of single or double spring-loaded pins (spring clips) on each foot piece to adjust, "]
[null,"Foot pieces to slide freely over the full extensibility and can be easily disassembled, 1mm minimum clearance between sliding parts,"]
[null,"Ferrule: removable non-slip fitted with steel insert, injection-moulded rubber PVC, "]
[null,"Double or single pin type spring clip at least 6mm diameter, electroplated to prevent corrosion"]
[null,"Must comply to the latest issue of CKS 112 for aluminium part"]
[null,"Each"]
["RT40-06-001","Wheels/castors"]
[null,"For Pulpit Walker, Reverce Walker, with wheels/castors"]
[null,"Wheels/castor size:  ±75mm in diameter, non-marking, rubberized"]
[null,"The castors must fit and be compatible with the items awarded for:\n1. Reverse Walker\n2. Rollator Gutter\n3. Pulpit Walker (for use by - Pediatric, Child, Child-Adult, and Adult)"]
[null,"Each"]
["RT40-07-001","Walking stick, T-handle"]
[null,"For adult use"]
[null,"Extruded anodised aluminium"]
[null,"Upper tube size:  ±25.4 x ±1.6mm"]
[null,"Lower tube size:  ±21.6 x ±1.4mm"]
[null,"The supplier should specify what size ferrule fit this crutch "]
[null,"Height adjustable: 650 - 930mm, "]
[null,"Base height: ±470mm"]
[null,"Extension height: ±500mm"]
[null,"The supplier should specify what size ferrule fit this crutch "]
[null,"Weight: ±600g"]
[null,"18 adjusting holes on the main tube for handle adjustable in 25mm increments which can suit various users. "]
[null,"Adjust by depression of two spring-loaded pins (spring clips), electroplated to prevent corrosion"]
[null,"T-handle (handgrip): injection-moulded grey propylene of grade 2340PC, immovable ergonomically designed"]
[null,"Spring loaded pins must be hard aluminium, brass or mild steel nickel plated and adjustment must be easy and smooth. "]
[null,"Ferrule/tip: Removable non-marking, non-slip ferrule with steel insert, injection-moulded  rubber PVC "]
[null,"Foot pieces to slide freely over the full extensibility and can be easily disassembled; max. clearance between sliding parts = 1mm,"]
[null,"Must comply to the latest issue of CKS 112"]
[]
["RT40-07-002","Walking Stick, Bent "]
[null,"For adult use"]
[null," AB quality polished batang rattan cane, shepherd’s-hook type with centre of bend over shaft of walking stick"]
[null,"Maximum Height; ±1040mm"]
[null,"Weight: ±750g"]
[null,"Ferrule/tip: Removable non-marking, non-slip ferrule with steel insert, injection-moulded  rubber PVC "]
[null,"Each "]
[null,"The supplier should specify what size ferrule fit this crutch "]
["RT40-07-003","Walking Stick, Tripod "]
[null,"For adult use"]
[null," Adjustable height between: 700mm- 930mm,"]
[null,"Base Height: ±500mm, "]
[null,"Extension height: ±500mm , "]
[null,"Weight: ±1.8kg weight,"]
[null,"The supplier should specify what size ferrule fit this crutch "]
[null,"Handgrip: injection-moulded propylene of grade 2340PC, immovable ergonomically designed"]
[null,"Adjust by depression of two spring-loaded pins (spring clips), double pin spring clip type at least 6mm in diameter electroplated to prevent corrosion, "]
[null,"3 legged steel base, "]
[null,"Steel tubing foot pieces to slide freely over the full extensibility and can be easily disassembled; max. clearance between sliding parts = 1mm, "]
[null,"Securely welded powder coated base, steel tubing of minimum diameter of 16.88mm and minimum 1.2mm wall thickness, uprights of steel tube - minimum 22mm diameter and minimum 1.5mm wall thickness"]
[null,"Ferrule/tip: removable non-slip, non-marking ferrules fitted with steel insert, injection-moulded rubber P.V.C ferrules, "]
[null,"Must comply to the latest issue of CKS 112 for aluminium part"]
[null,"Each"]
["RT40-07-004","Walking Stick, quadruped"]
[null,"For adult use"]
[null,"Extruded anodised aluminium"]
[null,"Adjustable height between: 700mm- 930mm,"]
[null,"Base Height: 500mm, "]
[null,"Extension height: 500mm , "]
[null,"The supplier should specify what size ferrule fit this crutch "]
[null,"Weight: ±1.8kg weight,"]
[null,"Handgrip: injection-moulded propylene of grade 2340PC, immovable ergonomically designed"]
[null,"Adjust by depression of two spring-loaded pins (spring clips), double pin spring clip type at least 6mm in diameter electroplated to prevent corrosion, "]
[null,"4 legged steel base, "]
[null,"Steel tubing foot pieces to slide freely over the full extensibility and can be easily disassembled; max. clearance between sliding parts = 1mm, "]
[null,"Securely welded powder coated base, steel tubing of minimum diameter of 16.88mm and minimum 1.2mm wall thickness, uprights of steel tube - minimum 22mm diameter and minimum 1.5mm wall thickness"]
[null,"18 adjusting holes on the base tube for handle adjustable in 25mm increments which can suit various users. "]
[null,"Ferrule/tip: removable non-slip, non-marking ferrules fitted with steel insert, injection-moulded rubber P.V.C ferrules, "]
[null,"Must comply to the latest issue of CKS 112 for aluminium part"]
[null,"Each                                                                                                                           "]
["RT40-07-005","Walking Stick, quadruped"]
[null,"Extruded anodised aluminium"]
[null,"Adjustable height between :700mm -1100mm"]
[null,"Base Height: ±600mm, "]
[null,"Extension height: ±600mm , "]
[null,"The supplier should specify what size ferrule fit this crutch "]
[null,"Weight: ±1.8kg weight,"]
[null,"Handgrip: injection-moulded propylene of grade 2340PC, immovable ergonomically designed"]
[null,"Adjust by depression of two spring-loaded pins (spring clips), double pin spring clip type at least 6mm in diameter electroplated to prevent corrosion, "]
[null,"4 legged steel base, "]
[null,"Steel tubing foot pieces to slide freely over the full extensibility and can be easily disassembled; max. clearance between sliding parts = 1mm, "]
[null,"Securely welded powder coated base, steel tubing of minimum diameter of 16.88mm and minimum 1.2mm wall thickness, uprights of steel tube - minimum 22mm diameter and minimum 1.5mm wall thickness"]
[null,"18 adjusting holes on the base tube for handle adjustable in 25mm increments which can suit various users. "]
[null,"Ferrule/tip: removable non-slip, non-marking ferrules fitted with steel insert, injection-moulded rubber P.V.C ferrules, "]
[null,"Must comply to the latest issue of CKS 112 for aluminium part"]
[null,"Each                                                                                                                           "]
["RT40-08-001","Heel walker, "]
[null,"Orthopaedic cast rubber, ribbed, wide flexible platform"]
[null,"Size: Small"]
[null,"Each"]
["RT40-08-002","Heel, walker"]
[null,"Orthopaedic cast rubber, ribbed, wide flexible platform"]
[null,"Size: Medium"]
[null,"Each"]
["RT40-08-003","Heel, walker"]
[null,"Orthopaedic cast rubber, ribbed, wide flexible platform"]
[null,"Size: Large"]
[null,"Each"]

{"sheet":"Sheet2","index":1,"n_rows":0,"n_cols":0,"rows_kept":0}

{"sheet":"Sheet3","index":2,"n_rows":0,"n_cols":0,"rows_kept":0}