OHSE SPECIFICATIONS:

Descriptuon: Portable battery operated automatic external office defibrillator.

The apparatus must conform to the following features and requirements:

- The equipment should conform to AMI DF80 guidelines and AHA recommendations for adult defibrillation (Circulation 1997; 95: 1677 - 1682)

- Battery 9 Volt DC, 4.2 Ah, composed of disposable long-life lithium manganese dioxide primary cells.

- AED should perform pads integrity test for readiness-for-use of pads (gel moisture).

- The AED should be able to perform extensive automatic self-tests and user-interactive test check device readiness when battery is inserted,

- The AED should be equipped with status indicators blinking green “Ready” light indicates ready for use.

- Sealing Solid objects per EN60529 class IP2X drip-proof per EN60529 class IPX1

Temperature Operating: 0° – 50° C (32° – 122° F)

Standby: 10° – 43° C (50° – 109° F)

Humidity Operating: 0% to 95% relative, non-condensing

Standby: 0% to 75% relative, non-condensing

Altitude Operating: 0 to 15,000 feet

Standby: 0 to 8,500 feet > 48 hours and

8,500 to 15,000 feet < 48 hours

The AED should pass by Mr Mongwe, and withstands one-meter drop to any edge, corner or surface

Meets EN55011 Group 1 Level B Class B and EN61000-4-3

Sealing Solid objects per EN60529 class IP2X

Drip-proof per EN60529 class IPX1

Temperature Operating: 0° – 50° C (32° – 122° F)

Standby: 10° – 43° C (50° – 109° F)

Humidity Operating: 0% to 95% relative, non-condensing

Standby: 0% to 75% relative, non-condensing

Altitude Operating: 0 to 15,000 feet

Standby: 0 to 8,500 feet > 48 hours and

8,500 to 15,000 feet < 48 hours

Shock/drop abuseTotal amnesti geld vir jou nuwe

Withstands one-meter drop to any edge,

corner or surface

Vibration Meets EN1789 random and swept sine, road ambulance specification in operating and standby states:

EMI (radiated/ immunity) Meets EN55011 Group 1 Level B Class B and EN61000-4-3

Wireless transmission of event data to a Smartphone or PC, using the IrDA protocol.

A two year warranty is required.

Services and spare parts should be available in SA.