* Supply and Installation of a Vertical Top Loading 50L Autoclave.

This autoclave MUST;

- conform to the South African National Standard SANS 347:2012,

- have the new certification of conformity to pressure requirements, design and fabrication according to the local SABS standard.

- Be installed by the accredited electrician.

Important Features:

Autoclave must have the following Safety Features:
1. Over pressure. Safety valve which can automatically open to release pressure.
2. Over temperature. The system that can switch power off automatically.
3. Door sensor. Sensor at door closing position, to check the door status. If door is open, operation must stop.
4. Anti-dry out function. If water level is too low, system must switch power off and displays error code on screen.
5. Over current. Over voltage or electric leakage protection apparatus-air switch.
6. Door cover. The door must be covered with safety shell made of thermal insulation material, which can protect operator from burning
by heat and looks nice.

Technical Features:
* Water inlet filter to keep high water quality.

* Optional printer for recording sterilization data.
* LED screen displays all steps of full process as well as error codes.
* Independent quick service window, easy for repairing and maintenance.
* Automatic selecting quick or low exhaust to avoid sterilization liquid overflowing.
* Sterilization programs for dressings, equipment, rubber, culture medium, waste, etc.
* Hand wheel type of quick-open door structure, self-expandable sealing ring, easy operating and safe.
* Complete process of filling water, dynamic pulse exhaust, heating, sterilizing, draining and drying runs automatically.

Specification

| Capacity | 50L |
| --- | --- |
| Chamber Size(mm) | φ386*514 |
| Chamber Material | SUS304 |
| Designed Pressure | 0.28Mpa |
| Designed Temp | 150℃ |
| Working Pressure | 0.22MPa |
| Working temp. | 105~136℃ |
| Temp. Display accuracy | ±0.1℃ |
| Timing range | 0~999min |
| Dynamic Pulse Exhaust Times | 0~9 times |
| Dynamic Pulse Exhaust Temp. | 105~134℃ |
| Dissolution Temp. Range | 60~100℃(programmable) |
| Temp. Holding Range | 40~60℃(programmable) |
| Consumption | 4.4KW |
| Power Supply | AC110/220V±10%, 50/60Hz |
| SS Basket | 1 pc |
| External Size(W*D*H)mm | 640*550*970 |
| Packing Size(W*D*H)mm | 750*745*1100 |
| Gross Weight(kg) | 102 |