Specifications for microplate photometer:

- Must read a 96 well microplate in less than 8 seconds.

- Must come with built in linear shaking facility.

- Must have a filter holder or filter wheel that can take at least 6 filters.

- Photometer must be supplied with at least the following filters: 405nm, 450nm, 492nm, 620nm.

- Must at least be compatible with wavelengths from 340nm to 850nm based on installation of different of different filters.

- Photometer must include all on-board software essential for operation.

- Power requirements: 240V, 50Hz, single phase.

- Light source: quartz halogen lamp.

- Dimensions: not larger than 41cm(D), 30cm(W), 24cm(H)

- User interface: on board and PC control (PC included).

- Resolution: at least 0.001 Abs.

- Linearity must be +/-2% or better for 0-3 ABS when reading a 96 well plate at 405nm

- Accuracy: Must be able to meet at least +/-1% for 0-3Abs or +/-0.003ABS (whichever is greater) at 405nm.

- Precision: CV must be </= 0.2% for 0.3 to 3 Abs at 405nm.

- Spares must be available for at least 10 years from date of commissioning.

- Service provider must have factory trained technicians and a fully equipped workshop. Training certificate/s for technicians must be supplied with quotations. Supplier must allow client to verify the existence of the workshop by means of a site visit.

- Quotation must include delivery, commissioning and training for two scientists or laboratory technicians.

- Photometer must come with a fitted surge protection plug.

|  | Requirement | Comply |  |
| --- | --- | --- | --- |
|  |  | Yes | No |
| 1 | Must read a 96 well microplate in less than 8 seconds. |  |  |
| 2 | Must come with built in linear shaking facility. |  |  |
| 3 | Must have a filter holder or filter wheel that can take at least 6 filters. |  |  |
| 4 | Photometer must be supplied with at least the following filters: 405nm, 450nm, 492nm, 620nm. |  |  |
| 5 | Must at least be compatible with wavelengths from 340nm to 850nm based on installation of different of different filters. |  |  |
| 6 | Photometer must include all on-board software essential for operation. |  |  |
| 7 | Power requirements: 240V, 50Hz, single phase. |  |  |
| 8 | Light source: quartz halogen lamp. |  |  |
| 9 | Dimensions: not larger than 41cm(D), 30cm(W), 24cm(H) |  |  |
| 10 | User interface: on board and PC control (PC included). |  |  |
| 11 | Resolution: at least 0.001 Abs. |  |  |
| 12 | Linearity must be +/-2% or better for 0-3 ABS when reading a 96 well plate at 405nm. |  |  |
| 13 | Accuracy: Must be able to meet at least +/-1% for 0-3Abs or +/-0.003ABS (whichever is greater) at 405nm. |  |  |
| 14 | Precision: CV must be </= 0.2% for 0.3 to 3 Abs at 405nm. |  |  |
| 15 | Spares must be available for at least 10 years from date of commissioning. |  |  |
| 16 | Service provider must have factory trained technicians and a fully equipped workshop. Training certificate/s for technicians must be supplied with quotations. Supplier must allow client to verify the existence of the workshop by means of a site visit. |  |  |
| 17 | Quotation must include delivery, commissioning and training for two scientists or laboratory technicians. |  |  |
| 18 | Photometer must come with a fitted surge protection plug. |  |  |