{"text_format":"3","workbook":"xlsx","n_sheets":1,"sheets":["GALLIUM FLUIDIC LABELLING MODUL"],"not_rendered":{"images":1,"defined_names":2},"read":"whole","date_system":"1900"}

{"sheet":"GALLIUM FLUIDIC LABELLING MODUL","index":0,"n_rows":34,"n_cols":6,"rows_kept":34,"merged":[[0,0,0,5],[1,0,1,5],[2,0,2,5],[3,0,3,5],[4,0,4,5],[5,0,5,5],[6,0,6,5],[7,0,7,5],[8,0,8,5],[9,0,9,5],[10,0,10,5],[11,0,11,5],[12,0,12,5],[13,0,13,5],[14,0,14,5]]}
["                              TENDER SPECIFICATION"]
["THE SUPPLY, DELIVERY, INSTALLATION, COMMISSIONING AND MAINTENANCE OF  GALLIUM FLUIDIC LABELLING MODULE FOR STEVE BIKO ACADEMIC HOSPITAL"]
["Note:                                                                                                                                                                                                                                                                                                                                                                                                                                                                                                                                                                                                                                                                                                   "]
["1. Technical Compliance"]
["The bidders must obtain a threshold of 2 points out of a total of 3 points  in respect of specifications(line) with a weight of 1."]
["2. Scoring of the bids"]
["2a. Weight: Each specification will have a weight of 1 or a hash (#), which indicates the level of importance, as follows:   "]
["1 -  A specification with a standard level of importance that must be complied with.,"]
["# -  Indicates a specification with a essential level of importance, where the bid offer will be disqualified, if the specification is not complied with."]
["2b. Score: The Bid Evaluation Committe will assign a score of 0 or1 for each specification, which indicates the level of compliance, as follows:"]
["0 -  The specification does not  comply.     1 -  The specification is fully compliant."]
["2c. Weighted and Total Score: The score will be multiplied by the weight to give the total score for each specification and be added up. "]
["2d. The bidders must comply as follows with the specification: "]
["i. The bidders must comply with all the specifications (line items) with a # weight. "]
[]
["A","B","C","D","E","F"]
["Item no.","Tender Specification","Weight","Comply (Yes/No)","Details of Bid Offer                                            (Provide your answers in this Colum. You are advised to be straight to the point)","Reference to Brochure / Technical Data Sheet (Provide page no for verification)"]
[1,"GENERAL DESCRIPTION :Ga-68 Fluidic Labeling Module (Compact synthesizer)"]
["1.1","Fully automated Ga-68 Fluidic Labeling Module","#"]
["1.2","Integratable to MUSA 68Ga hot cell","#"]
["1.3","Unit must be self shielded","#"]
["1.4","Compact and easily portable should not exceed the available working area of 22.5cm depth X 17.5cm width X 40cm height ","#"]
["1.5","Must be compatible with different dispensing systems including those produced by Comecer","#"]
["1.6","Must include filter integrity test","#"]
["1.7","Must be compatible with consumables from various suppliers","#"]
["1.8","Must be compatible with different Ga-68 generators including those produced by iThemba labs",1]
["1.0","Unit must provide a labeling yield exceeding 90 % ",1]
["1.10","Final product from elution must be less than 20 minutes",1]
["1.11","Systems software must be compatible to the Comecer dose calibrator","#"]
[]
[2,"Standard Consumables"]
["2.1","A minimum of 20 cassettes must be supplied ready to use in sealed pouches.","#"]
["2.2","Color coded Cassettes, Vials and reagent kits must be supplied","#"]
[null,"TOTAL  POINTS",3]