{"text_format":"3","workbook":"xls","n_sheets":2,"sheets":[" 3T specs"," 1.5T Specs"],"not_rendered":{"images":2,"external_links":1,"defined_names":4},"read":"whole","date_system":"1900","formulas":"cached_only"}

{"sheet":" 3T specs","index":0,"n_rows":318,"n_cols":5,"rows_kept":318,"merged":[[0,0,3,4],[4,0,4,4],[5,0,5,4],[6,0,6,4],[7,0,7,4],[8,0,8,4],[9,0,9,4],[10,0,10,4],[11,0,11,4],[12,0,12,4],[265,1,265,3]],"flags":["formulas_cached_only"]}
["TENDER SPECIFICATION FOR THE SUPPLY, DELIVERY, INSTALLATION, COMMISSIONING AND MAINTENANCE OF MRI MACHINES (3T) TO VARIOUS HOSPITALS IN GAUTENG "]
[]
[]
[]
["Note:                                                                                                                                                                                                                                                                                                                                                                                                                                                                                                                                                                                                                                                                                                   "]
["a. Weight: Each specification will have a weight of 1 or a hash (#), which indicates the level of importance, as follows:   "]
["1 -  A specification with a standard level of importance that must be complied with.     "]
["# -  Indicates a specification with essential level of importance, where the bid offer will be disqualified, if the specification is not complied with."]
["b. Score: The Bid Evaluation Committe will assign a score of 0 or 1 for each specification, as follows:"]
["0 -  The specification does not  comply.     1  -  The specification is fully compliant."]
["c. Weighted and Total Score: The score will be multiplied by the weight to give the total score for each specification and be added up. "]
["d. Bidders must comply as follows with the specification: "]
["i. The bidders must comply fully with the specifications marked with a # weight.                                                                                                                                                                                                                                                                                   ii. Bidders must obtain a total score of 76 or higher out of 105 points in respect of specifications (line) with a weight of 1."]
["A","B","C","D","E"]
["No.                                                                                                                                                                                                                                                          ","Tender specification","Weight","Comply (Yes/No)","Details of Bid Offer                                            (Provide your answers in this Colum. You are advised to be straight to the point)"]
["A"," Modality Details"]
["A 1","MAGNET"]
["A1.1","The Magnet Field Strength must be  3 Tesla active shielded","#"]
["A1.2","The Magnet shall be short bore cylindrical type and design","#"]
["A1.3","Bore length including all covers shall not exceed 180 cm",1]
["A1.4","State internal bore dimensions (L x W x H) not less than 70 cm",1]
["A1.4.1","State Maximum Vision Angle",1]
["A1.5","Magnet homogeneity shall meet the following specifications using the standard deviation VRMS (volume root-meansquare) using 24 plane plot measurement method:"]
["A1.5.1","50 cm DSV state",1]
["A1.5.2","48 cm DSV state",1]
["A1.5.3","45 cm DSV  state",1]
["A1.5.4","40 cm DSV state",1]
["A1.5.5","30 cm DSV state",1]
["A1.5.6","20 cm DSV state",1]
["A1.5.7","10 cm DSV  state",1]
["A1.6","Magnetic field stability shall be less than 0.1 ppm/hour or better",1]
["A1.7","The 5 Gauss/0.5mT fringe field shall be contained in an area of typically 2.4 m (radial) by 4.6m (axial).  State actual area",1]
["A1.8","State helium boil-off details, where applicable:"]
["A1.8.1","State typical cryogen refilling interval",1]
["A1.8.2","State maximum helium capacity.State latest technology/state helium capacity options /state latest cooling technology ",1]
["A1.8.3","State the weight of magnet with cryogens, table and covers (kg) in full operation ",1]
["A1.8.4","The bore diameter shall be a minimum of 70cm measured at the center in an operational mode","#"]
["A1.8.5","Min boil off rate should not be more than 0.02 1/hr under normal operating conditions","#"]
["A1.9","The tenderer must guarantee and ensure that there is no magnet vibration of the system, on the current location, with all possible imaging sequences. Give details.","#"]
["A1.9.1","State space required (Magnet, Electronics,UPSs and Control room). State type of magnet controllers-must be digital",1]
["B","MAGNET SAFETY"]
["B.1","Magnet must be equipped with quench exhaust leading to the outside of the building in event of Magnet quench to prevent injury to staff and patient/depending on type of cooling system available/state latest technology, vent pipe requirements state,if applicable",1]
["B.2","Magnet must be equipped with emergency ramp down unit for fast Magnetic Field reduction.","#"]
["B.3","Magnet must be equipped with ramp up procedure","#"]
["B.4","MRI Signage / Floor Sticker Warning Signs/ MRI Symbols Sign/ MRI Zone Signs Complete Set. “Do Not Enter” Floor Cones","#"]
["B.5","Detailed attachment explaining B.1 to B.4",1]
["C","GRADIENT SYSTEM"]
["C.1","Actively shielded hi-performance non-resonant gradient coil system is supplied.","#"]
["C.2","Gradient amplitude/peak strength must be at least 45mT/m measured per real axis plateau (100% duty cycle) not effective gradients: Sate options of highest gradients/Higher gradients will be an advantage.",1]
["C.3","Gradient duty cycle shall be 100%","#"]
["C.4","Minimum rise time from 0 to max. gradient amplitude shall be 240 micro-seconds or better",1]
["C.5","Gradient slew rate must be 200 m/mT/ms or more measured per real axis plateau, not effective values .State options of highest gradient",1]
["C.5.1","Shortest TE (SSh DWI, b=1000, 128 matrix)","#"]
["C.5.2","Shortest TE / TR (2D, 128 matrix) : 2D GRE","#"]
["C.5.3","Shortest TE / TR (2D, 128 matrix) : 3D GRE","#"]
["C.5.4","Shortest TE / TR (2D, 128 matrix) : 2D GRE EPI","#"]
["C.5.5","Shortest TE / TR (2D, 256 matrix) : 2D GRE ","#"]
["C.5.6","Shortest TE / TR (2D, 256 matrix) : 3D GRE","#"]
["C.5.7","Shortest TE / TR (2D, 256 matrix) : 2D GRE EPI","#"]
["C.5.8","Maximum scan matrix must be minimum 1024 or better",1]
["C.5.9","Maximum recon matrix must be stated",1]
["C.5.10","State Increment steps for matrix",1]
["C.5.11","State maximum number of slices",1]
["C.5.12","State minimum slice thickness (3D)",1]
["C.5.13","Gradient liniearity at 10 cm DSV","#"]
["C.5.14","Gradient linearity at 20 cm DSV","#"]
["C.5.15","Gradient linearity at 30 cm DSV","#"]
["C.5.16","Gradient linearity at 40 cm DSV","#"]
["C.5.17","Gradient linearity at 50 cm DSV","#"]
["C.5.18","The reconstruction speed at 100% FOV must not be less than 11000 reconstructions /sec.",1]
["C.6","Gradient upgrades to different levels should be possible without changing gradient coil. ",1]
["C.6.1","State high-end gradient specifications for optional gradient upgrade.",1]
["C.6.2","Max slew rate and amplitude and FOV at max slew rate and amplitude - FOV not less than 50cm.",1]
["C.6.3","Cooling for gradient coil; state type of coolant",1]
["C.7","The output linearity of the gradient amplifiers should be minimum  + 0.1% of peak",1]
["C.8","A child friendly scanning enviroment is required. Paediatric noise reduction accessories.State options",1]
["C.9","State advanced ergonomic features that characterize the configuration offered.",1]
["C.10","Noise should be kept to a minimum.Best acoustic noise reduction technology to be offered.State details",1]
["C.11","Free choice of flip angle while maintaining signal to noise ratio to be supported. Specify.",1]
["D","RADIOFREQUENCY SYSTEM"]
["D.1","Resonance frequency shall be 127MHz (3 Tesla) ",1]
["D.1.1","Direct Radio Frequency (RF) Transmitter System: Digital signal generation and processing.",1]
["D.1.2","The transmission system to be integrated in the magnet housing",1]
["D.1.3","Power output of transmitter amplifier rating shall be 35kW or better",1]
["D.1.4","State the bandwidth of the RF transmitter ",1]
["D.1.5","The system shall be equipped with RF fault protection limiting RF output in event of malfunction","#"]
["D.1.6","Frequency resolution of the RF synthesizer shall be 0.07 Hz /bit or more",1]
["D.2","The receive components shall be integrated into the magnet housing","#"]
["D.2.1","Phased array acquisition system with minimum of sixteen independent digital RF receiver channels or more, if applicable to the system. State commercial available channel upgrade path and commercial available coils utilizing the number of channels. Quote as optional items by the vendors.",1]
["D.2.2","The system should have at least 48 independent RF receiver channels or better, if applicable to the system",1]
["D.2.3","Phase resolution shall be 0.1 degree/bit or better",1]
["D.2.4","Noise of the preamplifier to be ≤ 0.5 decibels",1]
["D.3","Transmitter amplitude shall be 16 bit control or more",1]
["D.3.1","Can all of these coil elements be seamlessly used in one study without coil changes and without repositioning of the patient.",1]
["D.4","State maximum number of simultaneously connected coil elements. Preference will be given to highest number. Indicate coil combination.",1]
["D.4.1","Phase resolution shall be 0.1 degree/bit or more.",1]
["D.5","The Digital receiver should use at least 16 bit 1Mhz analogue to digital converter for   each receiver channel. State sampling method, sampling rate and bandwidth per channel. Receiver signal resolution should be 32 bit ",1]
["D.6","State number of RF cabinets required",1]
["E","RF COILS"]
["E.1","The bidder must supply the latest Integrated Coil Technology . Please state how the Coil Philosophy of the unit contributes and improve the Image Quality and Workflow. ",1]
["E.2","Integrated circular polarised / quadrature transmit and receive body coil.","#"]
["E.3","Minimum diameter of 70cm is  required . The bore to be measured at centre of the system and not on the flare.","#"]
["E.4","Connection of two or more phased array coils for simultaneous use.","#"]
["E.5","Connected coils must be detected automatically","#"]
["E.6","It must be possible to select active coil or elements from the main console","#"]
["E.7","Coil pre-amplifiers shall be on the patient table connector and coils shall be interchangeable and be of light construction.",1]
["E.8","Maximum coils connection to do full body scan. Give details of coils and connections (short cable coil connections preferred)","#"]
["E.9","Standared coils for the following applications should be available with  the system :"]
["E.9.1","Body (Chest, abdomen, pelvis).","#"]
["E.9.2","Tempo Mandibula Joint (TMJ) kit.  ",1]
["E.9.3","Head ","#"]
["E.9.4","Head-neck  ","#"]
["E.9.5","C/T/L spine  ( State what coil is used for imaging of the total spine cervical, thoracic and lumbar)","#"]
["E.9.6","Whole spine.","#"]
["E.9.7","Neck soft tissue.","#"]
["E.9.8","Extremities and joints: wrist, hips, shoulder,  knee,  foot/ankle,","#"]
["E.9.9","Breast . State solution for Breast MR with optional interventional capabilities. Quote interventional capabilities as optional.","#"]
["E.9.10","Prostate, colon, cervix. ","#"]
["E.9.11","Whole body vascular (Peripheral vascular array).","#"]
["E.9.12","Peripheral vascular.","#"]
["E.9.13","Neurovascular.","#"]
["E.9.14","Dedicated cardiac.","#"]
["E.9.15","Dedicated pediatric coils.","#"]
["E.9.16","Supply a cupboard in the room for storage of all coils provided or alternatively sufficient coil cabinets to house the coils.","#"]
["E.10","Other features:"]
["E.10.1","Coils to facilitate optimal facial, calvarial and intracranial imaging in the child from 1.5-60kgs, including visualisation of the orbits, optic nerves, optic tracts, pituitary fossa, hypothalamus, temporal lobe, posterior fossa and cerebral cortex. Coils to provide optimal imaging including, but not limited to the following clinical settings:",1]
["F","PATIENT SUPPORT/TABLE and MANAGEMENT/ PATIENT COMFORT"]
["F.1","Patient table must be dockable/detachable. Dockable patient transport system for simplified patient preparation, handling and transportation from preparation room to the MR scanner, without repositioning the patient.","#"]
["F.2","Two dockable tables must be included in the total price","#"]
["F.2.1","Two comfortable patient mattresses to be included","#"]
["F.3","Dual table control panels shall be located at either side of aperture/gantry for easy access. State if controls are available on the rear of the magnet","#"]
["F.4","Centering  laser light beams for anatomical references.",1]
["F.5","Table movement shall be controlled from both gantry and operator console","#"]
["F.6","Patient table shall be equipped with manual override for quick removal of patient from the magnet-bore in case of emergency","#"]
["F.7","Maximum load must be  minimum of 200kg or better. State actual maximum load.","#"]
["F.8","State maximum weight for vertical and horizontal table movement:",1]
["F.8.1","Vertical table movement:"]
["F.8.1.1","Minimum table height from floor",1]
["F.8.1.2","Maximum table height from floor",1]
["F.8.1.3","Maximum vertical speed",1]
["F.8.2","Horizontal table movement:"]
["F.8.2.1","Maximum horizontal range of table movement",1]
["F.8.2.2","Maximum speed:",1]
["F.8.2.3","Accuracy of repositioning",1]
["F.9","Physiologic measurement unit essential with display of ECG, respiration and pulse at the main console and gantry.","#"]
["F.10","Two way in-bore intercom system shall allow communication with patient while gradient is running","#"]
["F.11","Hand held alarm button for patient signalling","#"]
["F.12","Table restraining/immobilization straps are required per table","#"]
["F.13","Additional height adjustable patient transport stretcher –MR compatible max load 250kg to be supplied.",1]
["F.14","Paediatric friendly enviroment ( Bidder to provide details).",1]
["F.15","Integrated music for patient shall be included.",1]
["F.16","Variable patient lighting to be included.",1]
["F.17","State of the art technology is required to allow young patients to see outside the magnet during the examination. Give details",1]
["F.18","MRI compatible sand bags and sponges in various sizes and shapes.","#"]
["F.19","Ventilation in the gantry /Fresh air supply","#"]
["F.20","Auto-voice","#"]
["G","WORKSTATION:DOCTORS CONSOLE"]
["G.1","Two dedicated advanced post processing workstations will be supplied,  with extended memory.","#"]
["G.2","The dedicated workstation will be used for all post processing.  The system must include licenses for all 3D reconstruction and post processing requirements.  Please give details and options where applicable.","#"]
["G.3","Post processing on workstation to be compatible with software requested  on console.","#"]
["G.4","High quality read/write CD/DVD storage device required.","#"]
["G.5","DICOM compatible and configured to connect to a PACS system. Vendor must connect to the PACS system at no extra cost.","#"]
["H","OPERATOR USER INTERFACE"]
["H.1","One acquisition console with additional post processing workstation to be supplied.","#"]
["H.2","Two 23 inch or larger hi-resolution colour LCD flat-panel flicker-free monitor with undistorted image display required.","#"]
["H.3","Monitor 1024 x 1024 pixel resolution required or better","#"]
["H.4","User interface providing flexible multi-tasking in foreground or background (scanning, filming, reconstruction).","#"]
["H.5","Four TB external hard drive as storage devices required",1]
["H.6","All standard image processing features are required in addition the following advanced features must be available:"]
["H.6.1","Multi-planner reformatting (MPR).","#"]
["H.6.2","Multi-projection volume rendering (MPVR).","#"]
["H.6.3","Maximum intensity projection (MIP).","#"]
["H.6.4","MR Angiography processing.","#"]
["H.6.5","3D surface rendering.","#"]
["H.6.6","MR Hydrography MRCP/Urography/Myelography.","#"]
["H.6.7","Image add/subtract.","#"]
["H.6.8","CT image display.","#"]
["H.6.9","CT image integrating must be offered. ","#"]
["H.6.10","State any special features to improve productivity and consistency (e.g. automated scanning procedures) i.e  qualitative assessment.",1]
["I","ACQUISITION  PARAMATERS,  APPLICATION SOFTWARE  AND  IMAGING  TECHNIQUES"]
["I.1","Please explain the different techniques available to reduce motion artefacts to facilitate paediatric imaging.",1]
["I.2","Display slice and slab thickness and left and right designation:"]
["I.2.1","2D","#"]
["I.2.2","3D","#"]
["I.3","State number of slices:"]
["I.3.1","2D","#"]
["I.3.2","3D","#"]
["I.4","Variable field of view required to maximum of 48cm or better. State details",1]
["I.5","Standard/Conventional and fast imaging techniques/ Sequences required for all the following applications with dedicated post processing:"]
["a","State standard sequences in offered software package","#"]
["b","State maximum parallel imaging factor possible for offered software package (and compatible coils)","#"]
["c","State sequences that are compatible with parallel imaging","#"]
["d","List optional sequences and software packages, in addition to what is listed below. State details and costs.",1]
["I.5.1","Neuro:"]
["i","Brain","#"]
["ii","Conventional imaging sequences","#"]
["iii","Spectroscopy (Single voxel and multi voxel). State if compatible with parallel imaging 2D / 3D","#"]
["iv","Diffusion weighted","#"]
["v","Perfusion weighted. Sate sequences used.","#"]
["vi","Fast MRI sequences. Sate sequences used.","#"]
["vii","Tractography  / Fibre tracking ",1]
["viii","Spine","#"]
["I.5.2","Cardiac:"]
["i","Functional evaluation","#"]
["ii","Morphology","#"]
["iii","Valvular analysis","#"]
["iv","Coronary artery imaging","#"]
["v","Viability studies","#"]
["vi","ECG gating","#"]
["vii","Myocardial tagging","#"]
["I.5.3","Abdomen:"]
["i","Conventional sequences","#"]
["ii","Virtual endoscopy","#"]
["iii","MRCP","#"]
["iv","Small bowel studies","#"]
["v","Dynamic liver contrast studies","#"]
["I.5.4","Chest:"]
["i","Conventional sequences","#"]
["ii","Mammography","#"]
["iii","Dynamic (3D FSGRE etc.)","#"]
["iv","Biopsy capability","#"]
["v","Pulmonary ventilation studies","#"]
["I.5.5","Pelvis:"]
["i","Conventional sequences","#"]
["ii","Prostate imaging (including spectroscopy)","#"]
["iii","Dynamic pelvic floor imaging","#"]
["I.5.6","Vascular:"]
["i","TOF (2D/3D)","#"]
["ii","PC (2D/3D)","#"]
["iii","Contrast enhancement for:"]
["iv","Neck vessel studies","#"]
["v","Intracranial vessel studies","#"]
["vi","Aortic arch and branches","#"]
["vii","Abdominal aorta and outflow","#"]
["viii","Upper limbs","#"]
["ix","Pulmonary arteries","#"]
["x","Renal arteries","#"]
["xi","Multistep MR angiography","#"]
["I.5.7","Head and neck:"]
["i","Conventional sequences","#"]
["ii","IAM's","#"]
["iii","TMJ's","#"]
["J","PERFORMANCE EVALUATION PHANTOMS REQUIRED / QUALITY CONTROL"]
["J.1","Performance evaluation software and appropriate phantoms and quality assurance kits required for evaluation of image quality  as per licensing conditions.State details","#"]
["K","SAFETY"]
["K.1","The Magnet system shall include an emergency ramp down unit (RDU) for   fast reduction of the magnetic field.","#"]
["K.2","The Magnet system shall include an emergency ramp down unit (RDU) for   fast reduction of the magnetic field.",1]
["K.3","Real-time SAR calculation shall be performed by software to ensure that RF power   levels comply with regulatory guidelines and be displayed on each image.","#"]
["K.4","The system shall have manual override of the motor drive for quick removal of the   patient from the magnet bore.","#"]
["K.5","The system shall have RF fault protection hardware (RF transmit enable limit and   RF power and duty cycle limit) to limit The RF output in The event of equipment   malfunction.","#"]
["K.6","Pinpoint walk-through metal detector. Dedicated Ferro-magnetic detector with a pre-warning on approach well outside MRI room, not a conventional metal detector.  Highest sensitivity rating for the MRI environment.","#"]
["K.6.1","Hand held metal detector with different setting sensitivity.State details",1]
["K.6.2","The detector shall detect but not have any effect on heart pacemakers",1]
["K.7","A UPS System (Uninterrupted Power Supply) must be provided  for each workstation and should be able to run for at least 15 minutes. Must be included in the tender price.","#"]
["K.7.1","State details of kVA of UPS provided for the workstations.",1]
["K.8","A UPS System (Uninterrupted Power Supply) must be provided for the MRI Unit and  included in the tender price.State details of kVA of UPS provided for the unit.\nUPS must be 160kva or higher for the complete MRI system with sealed MF batteries for 30min back-up to support the RF amplifier and the gradients during a power failure. must be compatible with the MRI system,to include a 200kVa  isolation transformer.\n\n","#"]
["K.9","Give details of comprehensive magnet safety and MRI safety /include details of isolation transformer,earthing details,anti-theft cables,UPS and stability of power supply requirements","#"]
["K.10","Air-conditioning: Water chiller with filter connected to UPS","#"]
["L","ADDITIONAL  EQUIPMENT"]
["L.1","Extra patient monitoring screen at the console.State details",1]
["L.2","Please provide the specifications of the unit on offer .",1]
["L.3","MRI integrated pulse oximeter for adults and children. This unit must have probes for adult / paediatrics & neonates. ","#"]
["M","MR FLUROSCOPY (OPTIONAL)"]
["M.1","A MR Fluoroscopy mode with real-time acquisition and display techniques is required as an  option. Please State the soft and hardware required if this is optional.",1]
["M.2","State AI technologies available",1]
["N"," STANDARD ACCESSORIES"]
["N.1","MRI compatible accessories for Neonates,Paediatrics and Adults.State details.","#"]
["N.2","MRI compatible Incubator for Neonatal Imaging. ","#"]
["N.3","MRI compatible contrast injector interphased with MR unit. ","#"]
["N.4","200 syringes and 200 matching syringe lines.","#"]
["N.5","Drug Infusion pump compatible to MRI.","#"]
["N.6","List of standard accessories included but not specified in the tender.",1]
["N.7","Patient Monitoring Video Camera System. Direct connection to monitor. All parts in the MRI room must be MRI compatible","#"]
["O","TECHNOLOGY"]
["O.1","Innovation. Give explanation",1]
["O.2","Unique. Give explanation",1]
["O.3","Clinical benefits. Give detailed explanation",1]
["P","OPTIONAL ACCESSORIES"]
["P.1","The tenderer to give full description and pricing of optional accessories available for the system. Each item must be priced separately.Quote prices",1]
["P.2","MRI compatible laryngoscopes, a fully equipped MRI compatible emergency trolley and MRI compatible Stethoscope.",1]
["P.3","Scale weighing w/large dial type 2 150kg ",1]
["P.4","Patient slide for patient transfer 200g. (PVC Body Transfer Board and PVC Body Transfer Board Storage Brackets)",1]
["P.5","MRI compatible wheelchair with footrests",1]
["P.6","MRI Heavy Duty / Extra Wide Folding Walker.",1]
["P.7","MRI Transfer / Gait Belt.",1]
["P.8","MRI compatible steps / MRI Double Step Stool with Handrail",1]
["P.9","I.V. Pole, MRI compatible",1]
["P.10","Plastic chairs in magnet room",1]
["P.11","Chair operator high back swivel with arm rests. ",1]
["P.12","Chair draughtsman mid armrests",1]
["P.13","Shelve loose storage and display, wood.","#"]
["P.14","MRI Compatible Fire Extinguisher",1]
["P.15","Hoist mobile electric hydraulic.","#"]
["P.16","One scanner cum copier and color printer. (Fax/printer/copier/scanner)   ","#"]
["P.17","Dicom color printer / camera for 8\"x10\" paper/transparency color printing, 500 transparencies to be included.","#"]
["P.18","Modem for remote service support.",1]
["Q","MRI COMPATIBLE ANAESTHETIC UNIT   "]
["Q.1","MRI Compatible Anaesthetic unit  for General Anaesthetics of the paedeatric patients is required. Please state the specifications of the unit on offer. Please also note the warranty and customer support agreement terms. ","#"]
["Q.2","MRI compatible anaesthetic machine with ECG, NIBP, SPO2 & Gas  capnography   to suit adults, pediatric and neonates. ECG: Echo Cardiogram / NIBP: Non invasive BP  / SPO2: Oxygen saturation /   Gas Capnography: CO2 measurement","#"]
["Q.3","Anaesthesia machine must have a scavenger and the room must have an outlet for the scavenger for anaesthetic gases to be removed from the magnet room","#"]
["R","LASER CAMERA"]
["R.1","A state of the art dry laser camera to be included. Specify:  \nMake\nModel\n",1]
["R.2","UPS  for the dry laser camera to be included","#"]
["S","GENERAL ISSUES"]
["S.1","IT INTEGRATION "]
["S.1.1","The system must have the following DICOM 3 compatibility:"]
["S.1.1.1","DICOM Send / Receive","#"]
["S.1.1.2","DICOM Storage Commitment (SC):","#"]
["S.1.1.3","DICOM query/retrieve","#"]
["S.1.1.4","DICOM worklist","#"]
["S.1.1.5","DICOM Basic Print","#"]
["S.1.1.6","DICOM MPPS (Modality Performed Procedure Steps)","#"]
["S.2","The system must comply to the following IHE profiles (Integrating the Healthcare Enterprise)","#"]
["S.2.1","Scheduled Workflow (SWF)",1]
["S.2.2","Patient Information Reconciliation (PIR)",1]
["S.2.3","Simple Image and Numeric Reports (SINR)",1]
["S.2.4","System to be integrated to the existing PACS","#"]
[null,"TOTAL:",105]

{"sheet":" 1.5T Specs","index":1,"n_rows":317,"n_cols":6,"rows_kept":317,"merged":[[0,0,3,5],[4,0,4,5],[5,0,5,5],[6,0,6,5],[7,0,7,5],[8,0,8,5],[9,0,9,5],[10,0,10,5],[11,0,11,5],[12,0,12,5],[16,1,16,3],[38,1,38,3],[44,1,44,3],[77,1,77,3],[96,1,96,3],[124,1,124,3],[153,1,153,3],[160,1,160,3],[177,1,177,3],[241,1,241,3],[243,1,243,3],[257,1,257,3],[261,1,261,3],[264,1,264,3],[272,1,272,3],[276,1,276,3]],"flags":["formulas_cached_only"]}
["TENDER SPECIFICATION FOR THE SUPPLY, DELIVERY, INSTALLATION, COMMISSIONING AND MAINTENANCE OF  MRI MACHINES (1.5 T)  AT VARIOUS HOSPITALS IN GAUTENG"]
[]
[]
[]
["Note:                                                                                                                                                                                                                                                                                                                                                                                                                                                                                                                                                                                                                                                                                                   "]
["a. Weight: Each specification will have a weight of 1 or a hash (#), which indicates the level of importance, as follows:   "]
["1 -  A specification with a standard level of importance that must be complied with.     "]
["# -  Indicates a specification with  essential level of importance, where the bid offer will be disqualified, if the specification is not complied with."]
["b. Score: The Bid Evaluation Committe will assign a score of 0 or 1, for each specification, as follows:"]
["0 -  The specification does not  comply.     1 -  The specification is fully compliant.           "]
["c. Weighted and Total Score: The score will be multiplied by the weight to give the total score for each specification and be added up. "]
["d.The bidders must comply as follows with the specification: "]
["i. Bidders must comply fully with the specifications marked with a # weight.                                                                                                                                                                                                                               ii.Bidders must obtain a total score of 50 or higher out of 102 points in respect of specifications (line) with a weight of 1."]
["A","B","C","D","E","F"]
[null,"Tender Specification","Weight","Comply (Yes/No)","Details of Bid Offer                                            (Provide your answers in this Colum. You are advised to be straight to the point)","Reference to Brochure / Technical Data Sheet and provide page numbers for verification"]
["A"," Modality Details"]
["A 1","MAGNET"]
["A1.1","The Magnet Field Strength shall be at least 1.5 Tesla active shielded.","#"]
["A1.2","The Magnet must be short closed bore cylindrical type and design.","#"]
["A1.3","Bore length including all covers shall not exceed 180 cm",1]
["A1.4","State internal bore dimensions (L x W x H) not less than 70cm inner magnet.",1]
["A1.5","Magnet homogeneity should meet the following specifications using the standard deviation VRMS (volume root-meansquare) using 24 plane plot measurement method:"]
["A1.5.1","50 cm DSV (Diameter Spherical Volume) state (at most 5ppm)",1]
["A1.5.2","45 cm DSV  state (at most 1.2 ppm)",1]
["A1.5.3","40 cm DSV state (at most 1 ppm)",1]
["A1.5.4","30 cm DSV state (at most 0.5 ppm)",1]
["A1.5.5","20 cm DSV state (at most 0.1 ppm)",1]
["A1.5.6","10 cm DSV  state (at most 0.02 ppm)",1]
["A 1.6","Magnetic field stability shall be less than 0.1 ppm/hour or better",1]
["A 1.7","The 5 Gauss/0.5mT fringe field must be contained in an area of typically 2.4 m (radial) by 4.0m (axial).  State actual area.",1]
["A 1.8","State helium boil-off details, where applicable:"]
["A1.8.1","State typical cryogen refilling interval",1]
["A1.8.2","State maximum helium capacity.State latest technology/state helium capacity options /state latest cooling technology ",1]
["A1.8.3","State the weight of magnet with cryogens, table and covers (kg) in full operation ",1]
["A1.8.4","The bore diameter shall be a minimum of 70cm measured at the center in an operational mode","#"]
["A1.8.5","Min boil off rate should not be more than 0.02 1/hr under normal operating conditions","#"]
["A1.9","The tenderer must guarantee and ensure that there is no magnet vibration of the system, on the current location, with all possible imaging sequences. Give details.","#"]
["A1.9.1","State space required (Magnet, Electronics,UPSs and Control room). State type of magnet controllers-must be digital",1]
["B.","MAGNET SAFETY"]
["B.1","Magnet must be equipped with quench exhaust leading to the outside of the building in event of Magnet quench to prevent injury to staff and patient/depending on type of cooling system available/state latest technology, vent pipe requirements state,If applicable",1]
["B.2","Magnet must be equipped with emergency ramp down unit for fast Magnetic reduction, state details.","#"]
["B.3","Magnet must be equipped with ramp up procedure","#"]
["B.4","MRI Signage / Floor Sticker Warning Signs/ MRI Symbols Sign/ MRI Zone Signs Complete Set. “Do Not Enter” Floor Cones","#"]
["B.5","Detailed attachment explaining B.1 to B.4",1]
["C","GRADIENT SYSTEM (Provide the state of the art and most recent technology details)"]
["C.1","Actively shielded hi-performance non-resonant gradient coil system is supplied.","#"]
["C.2","Gradient amplitude/peak strength must be at least 33mT/m measured per real axis plateau (100% duty cycle) not effective gradients: Sate options of highest gradients/Higher gradients will be an advantage.",1]
["C.3","Gradient duty cycle shall be 100%.","#"]
["C.4","Gradient rise time from 0 to max gradient amplitude must be 240 micro-seconds or more.",1]
["C.5","Gradient slew rate must be 120 m/mT/ms or more measured per real axis plateau, not effective values .State options of highest gradient",1]
["C.5.1","Shortest TE (SSh DWI, b=1000, 128 matrix)","#"]
["C.5.2","Shortest TE / TR (2D, 128 matrix) : 2D GRE","#"]
["C.5.3","Shortest TE / TR (2D, 128 matrix) : 3D GRE","#"]
["C.5.4","Shortest TE / TR (2D, 128 matrix) : 2D GRE EPI","#"]
["C.5.5","Shortest TE / TR (2D, 256 matrix) : 2D GRE ","#"]
["C.5.6","Shortest TE / TR (2D, 256 matrix) : 3D GRE","#"]
["C.5.7","Shortest TE / TR (2D, 256 matrix) : 2D GRE EPI","#"]
["C.5.8","Maximum scan matrix must be minimum 1024 or better",1]
["C.5.9","Maximum recon matrix must be stated",1]
["C.5.10","State Increment steps for matrix","#"]
["C.5.11","State Maximum number of slices",1]
["C.5.12","State Minimum slice thickness (3D)",1]
["C.5.13","Gradient linearity at 20 cm DSV","#"]
["C.5.14","Gradient linearity at 50 cm DSV","#"]
["C.5.15","Gradient linearity at 30 cm DSV","#"]
["C.5.16","Gradient linearity at 40 cm DSV","#"]
["C.5.17","Gradient linearity at 50 cm DSV","#"]
["C.5.18","The reconstruction speed at 100% FOV must not be less than 11000 reconstructions /sec.",1]
["C.6","Gradient upgrades to different levels should be possible without changing gradient coil. ",1]
["C.6.1","State high-end gradient specifications for optional gradient upgrade.",1]
["C.6.2","Max slew rate and amplitude and FOV at max slew rate and amplitude - FOV not less than 50cm.",1]
["C.6.3","Cooling for gradient coil; state type of coolant.",1]
["C.7","The output linearity of the gradient amplifiers should be minimum of  + 0.1% of peak.",1]
["C.8","A child friendly scanning enviroment is required. Paediatric noise reduction accessories.State options",1]
["C.9","State advanced ergonomic features that characterize the configuration offered.",1]
["C.10","Noise should be kept to a minimum.",1]
["C.11","Free choice of flip angle while maintaining signal to noise ratio to be supported. Specify.",1]
["D.","RADIOFREQUENCY SYSTEM"]
["D.1","Resonance frequency must be 63.86MHz (1.5 Tesla). ",1]
["D.1.1","Direct Radio Frequency (RF) Transmitter System: Digital signal generation and processing.",1]
["D.1.2","The transmission system to be integrated in the magnet housing.",1]
["D.1.3","Power output of transmitter amplifier rating must be 18kW or better",1]
["D.1.4","State the bandwidth of the RF transmitter. ",1]
["D.1.5","The system must be equipped with RF fault protection limiting RF output in event of malfunction.","#"]
["D.1.6","Frequency resolution of the RF synthesizer shall be 0.07 Hz /bit or more",1]
["D.2","The receive components shall be integrated into the magnet housing","#"]
["D.2.1","Phased array acquisition system with minimum of sixteen independent digital RF receiver channels or more, if applicable to the system. State commercial available channel upgrade path and commercial available coils utilizing the number of channels. Quote as optional items by the vendors.",1]
["D.2.2","The system should have at least 48 independent RF receiver channels or better, if applicable to the system",1]
["D.2.3","Phase resolution shall be 0.1 degree/bit or better",1]
["D.2.4","Noise of the preamplifier to be ≤ 0.5 decibels",1]
["D.3","Transmitter amplitude shall be 16 bit control or more",1]
["D.3.1","Can all of these coil elements be seamlessly used in one study without coil changes and without repositioning of the patient.",1]
["D.4","State maximum number of simultaneously connected coil elements. Preference will be given to highest number. Indicate coil combination.",1]
["D.4.1","Phase resolution shall be 0.1 degree/bit or more.",1]
["D.5","The Digital receiver should use at least 16 bit 1Mhz analogue to digital converter for   each receiver channel. State sampling method, sampling rate and bandwidth per channel.",1]
["D.6","State number of RF cabinets required",1]
["E.","RF COILS"]
["E.1","The bidder must supply the latest Integrated Coil Technology. Please state how the Coil Philosophy of the unit contributes and improve the Image Quality and Workflow. ",1]
["E.2","Integrated circular polarised / quadrature transmit and receive body coil.","#"]
["E.3","Minimum diameter of 70cm is  required . The bore to be measured at centre of the system and not on the flare.","#"]
["E.4","Connection of two or more phased array coils for simultaneous use.","#"]
["E.5","Connected coils must be detected automatically.","#"]
["E.6","It must be possible to select active coil or elements from the main console.","#"]
["E.7","Coil pre-amplifiers shall be on the patient table connector and coils shall be interchangeable and be of light construction.",1]
["E.8","Maximum coils connection to do full body scan. Give details of coils and connections (short cable coil connections preferred)","#"]
["E.9","Standared coils for the following applications should be available with  the system :"]
["E.9.1","Body (Chest, abdomen, pelvis).","#"]
["E.9.2","Tempo Mandibula Joint (TMJ) kit.  ",1]
["E.9.3","Head ","#"]
["E.9.4","Head-neck  ","#"]
["E.9.5","C/T/L spine  ( State what coil is used for imaging of the total spine cervical, thoracic and lumbar)","#"]
["E.9.6","Whole spine.","#"]
["E.9.7","Neck soft tissue.","#"]
["E.9.8","Extremities and joints: wrist, hips, shoulder,  knee,  foot/ankle,","#"]
["E.9.9","Breast . State solution for Breast MR with optional interventional capabilities. Quote interventional capabilities as optional.","#"]
["E.9.10","Prostate, colon, cervix. ","#"]
["E.9.11","Whole body vascular (Peripheral vascular array).","#"]
["E.9.12","Peripheral vascular.","#"]
["E.9.13","Neurovascular.","#"]
["E.9.14","Dedicated cardiac.","#"]
["E.9.15","Dedicated pediatric coils.","#"]
["E.9.16","Supply a cupboard in the room for storage of all coils provided or alternatively sufficient coil cabinets to house the coils.","#"]
["E.10","Other features:"]
["E.10.1","Coils to facilitate optimal facial, calvarial and intracranial imaging in the child from 1.5-60kgs, including visualisation of the orbits, optic nerves, optic tracts, pituitary fossa, hypothalamus, temporal lobe, posterior fossa and cerebral cortex. Coils to provide optimal imaging including, but not limited to the clinical settings.","#"]
["F","PATIENT SUPPORT/TABLE and MANAGEMENT/ PATIENT COMFORT"]
["F.1","Patient table must be dockable/detachable. Dockable patient transport system for simplified patient preparation, handling and transportation from preparation room to the MR scanner, without repositioning the patient.","#"]
["F.2","Two dockable tables must be included in the total price","#"]
["F.3","Dual table control panels shall be located at either side of aperture/gantry for easy access. State if controls are available on the rear of the magnet.","#"]
["F.4","Centering  laser light beams for anatomical references.",1]
["F.5","Table movement shall be controlled from both gantry and operator console.","#"]
["F.6","Patient table shall be equipped with manual override for quick removal of patient from the magnet-bore in case of emergency.","#"]
["F.7","Maximum load must be  minimum of 200kg or better. State actual maximum load.","#"]
["F.8","maximum weight for vertical and horizontal table movement:"]
["F.8.1","Vertical table movement:"]
["F.8.1.1","Minimum table height from floor",1]
["F.8.1.2","Maximum table height from floor",1]
["F.8.1.3","Maximum vertical speed",1]
["F.8.2","Horizontal table movement:"]
["F.8.2.1","Maximum horizontal range of table movement",1]
["F.8.2.2","Maximum speed:",1]
["F.8.2.3","Accuracy of repositioning",1]
["F.9","Physiologic measurement unit essential with display of ECG, respiration and pulse at the main console and gantry.","#"]
["F.10","Two way in-bore intercom system shall allow communication with patient while gradient is running","#"]
["F.11","Hand held alarm button for patient signalling","#"]
["F.12","Table restraining/immobilization straps are required per table","#"]
["F.13","Additional height adjustable patient transport stretcher –MR compatible max load 250kg to be supplied.",1]
["F.14","Paediatric friendly enviroment ( Bidder to provide details).",1]
["F.15","Integrated music for patient shall be included.",1]
["F.16","Variable patient lighting to be included.",1]
["F.17","State of the art technology is required to allow young patients to see outside the magnet during the examination. Give details",1]
["F.18","MRI compatible sand bags and sponges in various sizes and shapes.","#"]
["F.19","Ventilation in the gantry /Fresh air supply","#"]
["F.20","Auto-voice","#"]
["G"," WORKSTATION: DOCTORS CONSOLE"]
["G.1","Two dedicated advanced post processing workstations will be supplied,  with extended memory.","#"]
["G.2","The dedicated workstation will be used for all post processing.  The system must include licenses for all 3D reconstruction and post processing requirements.  Please give details and options where applicable.","#"]
["G.3","Post processing on workstation to be compatible with software requested  on console.","#"]
["G.4","High quality read/write CD/DVD storage device required.","#"]
["G.5","DICOM compatible and configured to connect to a PACS system. Vendor must connect to the PACS system at no extra cost.","#"]
["G.6","DICOM compatible and configured to connect to a PACS system. Vendor must connect to the PACS system at no extra cost.","#"]
["H","OPERATOR USER INTERFACE"]
["H.1","One acquisition console with additional post processing workstation to be supplied.","#"]
["H.2","Two 23 inch or larger hi-resolution colour LCD flat-panel flicker-free monitor with undistorted image display required.","#"]
["H.3","Monitor 1024 x 1024 pixel resolution required or better","#"]
["H.4","User interface providing flexible multi-tasking in foreground or background (scanning, filming, reconstruction).","#"]
["H.5","Four TB external hard drive as storage devices required",1]
["H.6","All standard image processing features are required in addition the following advanced features must be available:"]
["H.6.1","Multi-planner reformatting (MPR).","#"]
["H.6.2","Multi-projection volume rendering (MPVR).","#"]
["H.6.3","Maximum intensity projection (MIP).","#"]
["H.6.4","MR Angiography processing.","#"]
["H.6.5","3D surface rendering.","#"]
["H.6.6","MR Hydrography MRCP/Urography/Myelography.","#"]
["H.6.7","Image add/subtract.","#"]
["H.6.8","CT image display.","#"]
["H.6.9","CT image integrating must be offered. ","#"]
["H.6.10","State any special features to improve productivity and consistency (e.g. automated scanning procedures) i.e  qualitative assessment.",1]
["I","ACQUISITION  PARAMATERS,  APPLICATION SOFTWARE  AND  IMAGING  TECHNIQUES "]
["I.1","Please explain the different techniques available to reduce motion artefacts to facilitate paediatric imaging.",1]
["I.2","Display slice and slab thickness and left and right designation:"]
["I.2.1","2D.","#"]
["I.2.2","3D.","#"]
["I.3","Display number of slices:"]
["I.3.1","2D.","#"]
["I.3.2","3D.","#"]
["I.4","Variable field of view required to maximum of 48cm or better. State.","#"]
["I.5","Standard/Conventional and fast imaging techniques/ Sequences required for all the following applications with dedicated post processing:"]
["a","List standard sequences in offered software package.","#"]
["b","List maximum parallel imaging factor possible for offered software package (and compatible coils).","#"]
["c","List sequences that are compatible with parallel imaging.","#"]
["d","List optional sequences and software packages, in addition to what is listed below. State details and costs.","#"]
["I.5.1","Neuro:"]
["i","Brain.","#"]
["ii","Conventional imaging sequences.","#"]
["iii","Spectroscopy (Single voxel and multi voxel). State if compatible with parallel imaging 2D / 3D.","#"]
["iv","Diffusion weighted.","#"]
["v","Perfusion weighted. Sate sequences used.","#"]
["vi","Fast MRI sequences. Sate sequences used.","#"]
["vii","Tractography  / Fibre tracking .","#"]
["viii","Spine.","#"]
["I.5.2","Cardiac:"]
["i","Functional evaluation.","#"]
["ii","Morphology.","#"]
["iii","Valvular analysis.","#"]
["iv","Coronary artery imaging.","#"]
["v","Viability studies.","#"]
["vi","ECG gating.","#"]
["vii","Myocardial tagging.","#"]
["I.5.3","Abdomen:"]
["i","Conventional sequences.","#"]
["ii","Virtual endoscopy.","#"]
["iii","MRCP.","#"]
["iv","Small bowel studies.","#"]
["v","Dynamic liver contrast studies & Liver DWI.","#"]
["I.5.4","Chest:"]
["i","Conventional sequences.","#"]
["ii","Mammography .","#"]
["iii","Dynamic (3D FSGRE etc.).","#"]
["iv","Biopsy capability.","#"]
["v","Pulmonary ventilation studies.","#"]
["I.5.5","Pelvis:"]
["i","Conventional sequences.","#"]
["ii","Prostate imaging (including spectroscopy).","#"]
["iii","Dynamic pelvic floor imaging.","#"]
["I.5.6","Vascular:"]
["i","TOF (2D/3D).","#"]
["ii","PC (2D/3D).","#"]
["iii","Contrast enhancement for:"]
["iv","Neck vessel studies.","#"]
["v","Intracranial vessel studies.","#"]
["vi","Aortic arch and branches.","#"]
["vii","Abdominal aorta and outflow.","#"]
["viii","Upper  & Lower limbs.","#"]
["ix","Pulmonary arteries.","#"]
["x","Renal arteries.","#"]
["xi","Multistep MR angiography.","#"]
["I.5.7","Head and neck:"]
["i","Conventional sequences.","#"]
["ii","IAM's (Internal Auditory Meatus).","#"]
["iii","TMJ's.","#"]
["I.6","Explain what techniques are available to reduce breathhold time.",1]
["J","PERFORMANCE EVALUATION PHANTOMS REQUIRED / QUALITY CONTROL"]
["J.1","Performance evaluation software and appropriate phantoms and quality assurance kits required for evaluation of image quality  as per licensing conditions.State details","#"]
["K","SAFETY"]
["K.1","The Magnet system shall include an emergency ramp down unit (RDU) for   fast reduction of the magnetic field.","#"]
["K.2","The Magnet system shall include an emergency ramp down unit (RDU) for   fast reduction of the magnetic field.",1]
["K.3","Real-time SAR calculation shall be performed by software to ensure that RF power   levels comply with regulatory guidelines and be displayed on each image.","#"]
["K.4","The system shall have manual override of the motor drive for quick removal of the   patient from the magnet bore.","#"]
["K.5","The system shall have RF fault protection hardware (RF transmit enable limit and   RF power and duty cycle limit) to limit The RF output in The event of equipment   malfunction.","#"]
["K.6","Pinpoint walk-through metal detector. Dedicated Ferro-magnetic detector with a pre-warning on approach well outside MRI room, not a conventional metal detector.  Highest sensitivity rating for the MRI environment.","#"]
["K.6.1","Hand held metal detector with different setting sensitivity.State details",1]
["K.6.2","The detector shall detect but not have any effect on heart pacemakers",1]
["K.7","A UPS System (Uninterrupted Power Supply) must be provided  for each workstation provided and  included in the tender price. ","#"]
["K.7.1","State details of kVA of UPS provided for the workstations.",1]
["K.8","A UPS System (Uninterrupted Power Supply) must be provided for the MRI Unit and  included in the tender price.State details of kVA of UPS provided for the unit.\nUPS must be 160kva or higher for the complete MRI system with sealed MF batteries for 30min back-up to support the RF amplifier and the gradients during a power failure. must be compatible with the MRI system,to include a 200kVa  isolation transformer.\n\n","#"]
["K.9","Give details of comprehensive magnet safety and MRI safety /include details of isolation transformer,earthing details,anti-theft cables,UPS and stability of power supply requirements","#"]
["K.10","Air-conditioning: Water chiller with filter connected to UPS","#"]
["L","ADDITIONAL  EQUIPMENT"]
["L.1","Extra patient monitoring screen at the console.State details",1]
["L.2","Please provide the specifications of the unit on offer .",1]
["L.3","MRI integrated pulse oximeter for adults and children. This unit must have probes for adult / paediatrics & neonates. ","#"]
["M","MR FLUROSCOPY (OPTIONAL)"]
["M.1","a  MR Fluoroscopy mode with real-time acquisition and display techniques is required as an  option. Please State the soft and hardware required if this is optional.",1]
["M.2","State AI technologies available",1]
["N"," STANDARD ACCESSORIES"]
["N.1","MRI compatible accessories for Neonates,Paediatrics and Adults.State details.","#"]
["N.2","MRI compatible Incubator for Neonatal Imaging. ","#"]
["N.3","MRI compatible contrast injector interphased with MR unit. ","#"]
["N.4","200 syringes and 200 matching syringe lines.","#"]
["N.5","Drug Infusion pump compatible to MRI.","#"]
["N.6","List of standard accessories included but not specified in the tender.",1]
["N.7","Patient Monitoring Video Camera System. Direct connection to monitor. All parts in the MRI room must be MRI compatible","#"]
["O","TECHNOLOGY"]
["O.1","Innovation. Give explanation",1]
["O.2","Unique. Give explanation",1]
["O.3","Clinical benefits. Give detailed explanation",1]
["P","OPTIONAL ACCESSORIES"]
["P.1","The tenderer to give full description and pricing of optional accessories available for the system. Each item must be priced separately.Quote prices",1]
["P.2","MRI compatible laryngoscopes, a fully equipped MRI compatible emergency trolley and MRI compatible Stethoscope.",1]
["P.3","Scale weighing w/large dial type 2 150kg ",1]
["P.4","Patient slide for patient transfer 200g. (PVC Body Transfer Board and PVC Body Transfer Board Storage Brackets)",1]
["P.5","MRI compatible wheelchair with footrests",1]
["P.6","MRI Heavy Duty / Extra Wide Folding Walker.",1]
["P.7","MRI Transfer / Gait Belt.",1]
["P.8","MRI compatible steps / MRI Double Step Stool with Handrail",1]
["P.9","I.V. Pole, MRI compatible",1]
["P.10","Plastic chairs in magnet room",1]
["P.11","Chair operator high back swivel with arm rests. ",1]
["P.12","Chair draughtsman mid armrests",1]
["P.13","Shelve loose storage and display, wood.",1]
["P.14","MRI Compatible Fire Extinguisher",1]
["P.15","Hoist mobile electric hydraulic.",1]
["P.16","One scanner cum copier and color printer. (Fax/printer/copier/scanner)   ",1]
["P.17","Dicom color printer / camera for 8\"x10\" paper/transparency color printing, 500 transparencies to be included.",1]
["P.18","Modem for remote service support.",1]
["Q","MRI COMPATIBLE ANAESTHETIC UNIT "]
["Q.1","MRI Compatible Anaesthetic unit  for General Anaesthetics of the paedeatric patients is required. Please state the specifications of the unit on offer. Please also note the warranty and customer support agreement terms. ","#"]
["Q.2","MRI compatible anaesthetic machine with ECG, NIBP, SPO2 & Gas  capnography   to suit adults, pediatric and neonates. ECG: Echo Cardiogram / NIBP: Non invasive BP  / SPO2: Oxygen saturation /   Gas Capnography: CO2 measurement","#"]
["Q.3","Anaesthesia machine must have a scavenger and the room must have an outlet for the scavenger for anaesthetic gases to be removed from the magnet room","#"]
["R","LASER CAMERA"]
["R.1","A state of the art dry laser camera to be included. Specify:  \nMake\nModel\n",1]
["R.2","UPS  for the dry laser camera to be included","#"]
["S","GENERAL ISSUES"]
["S.1","IT INTEGRATION "]
["S.1.1","The system must have the following DICOM 3 compatibility:"]
["S.1.1.1","DICOM Send / Receive","#"]
["S.1.1.2","DICOM Storage Commitment (SC):","#"]
["S.1.1.3","DICOM query/retrieve","#"]
["S.1.1.4","DICOM worklist","#"]
["S.1.1.5","DICOM Basic Print","#"]
["S.1.1.6","DICOM MPPS (Modality Performed Procedure Steps)","#"]
["S.2","The system must comply to the following IHE profiles (Integrating the Healthcare Enterprise)","#"]
["S.2.1","Scheduled Workflow (SWF)",1]
["S.2.2","Patient Information Reconciliation (PIR)",1]
["S.2.3","Simple Image and Numeric Reports (SINR)",1]
["S.2.4","System to be integrated to the existing PACS","#"]
[null,"TOTAL:",102]