{"text_format":"3","workbook":"xlsx","n_sheets":8,"sheets":["BP MACHINE","HB Meter","Multiparameter","Dual GLu &HbA1c","HbA1c","Glucometer","Cholesterol","Ketones"],"read":"whole","date_system":"1900"}

{"sheet":"BP MACHINE","index":0,"n_rows":70,"n_cols":4,"rows_kept":70,"merged":[[4,1,4,3],[5,1,5,3],[6,1,6,3]]}
[null,"Portable Upper Arm,Adult Blood Pressure Monitor, Electronic, NIBP","Complies Yes/No","Provide your answers /Comments in this Column.  You are advised to be straight to the point."]
["1","Description"]
[null,"Portable Electronic Upperarm blood pressure measuring device for adults"]
["2","Product Offered:"]
["2.1","Make:"]
["2.2","Model:"]
["2.3","Initial Manufacturing year of Model:"]
["3","Technical Specification"]
["3.1","The unit must measure blood pressure non-invasively."]
["3.2","Self test procedure at switch-on"]
["3.3","Must be suitable for adults 18 years and older"]
["3.4","The unit must use the oscillometric method of measurement."]
["3.5","Pressure measurements:"]
["3.6","Adult: Systolic: 30 - 270 mm Hg, Diastolic: 20 - 240 mm Hg, colour coding to be considered"]
["3.7","Auto zero when unit is switched on"]
["3.8","Over pressure protection: Automatic deflation at 300 mm Hg adult (error to be displayed in LCD)"]
["3.9","Average measurement time: approximately 30 s or less"]
["3.10","If measurement exceeds two minutes the cuff must deflate, an alarm must sound and likely faults indicated."]
["3.11","The air hose must be at least 1 metre long. "]
["3.12","LCD Display:"]
["3.12.1","The display must offer clear viewing under all lighting conditions , (display to include backlight)"]
["3.12.2","The following must be displayed:"]
["3.12.2a"," Systolic, diastolic pressures"]
["3.12.2b","Visual indication(colour codes) for lower, normal, high and critically high blood pressure"]
["3.12.2c","   Battery status"]
["3.14.2d","   Indication of mains power or battery operation"]
["3.13","ALARMS: "]
["3.13.1","All alarms must have default settings"]
["3.13.2","Violation of limits must result in both audible and/ visual alarms and indication of the parameters affected.  "]
["3.13.3","Low battery: Alarm or visual indication at least 10 minutes before battery power is exhausted"]
["3.14","Other system requirements"]
["3.14.1","The unit must operate from 220 V, 50 Hz mains supply adapter"]
["3.14.2","Adapter cable at least two metres long with SABS approved 15A three prong plug. "]
["3.14.3","The unit must include an internal lithium rechargeable battery. "]
["3.14.4","The unit must be able to function from the mains supply regardless of the battery condition."]
["3.14.5","Rechargeable battery must have the capacity for the unit to operate fully for at least five hours"]
["3.14.6","The rechargeable battery must be charged while the unit is operating from the mains supply (lithium batteries)."]
["3.14.7","The battery must automatically take over in the event of a mains power failure."]
["3.14.8","The case of the unit must be robust and impact resistant."]
["3.14.9","The design must be such that spilled liquids do not gain access to the internal working of the unit."]
["3.14.10","The monitor must be compact and stable to operate on the table or handheld."]
["3.14.11","The unit must be handed over in full operating order/fully functional"]
["3.14.12","A starter-pack of consumables must be supplied with the unit. "]
["4"," Standard Consumables"]
["4.1"," 1x  Reusable latex-free small adult cuff"]
["4.2"," 2 x Reusable  latex-free standard adult cuff"]
["4.3"," 1 xReusable  latex-free obese cuff"]
["4.4"," 1 x NIBP hose"]
["4.5"," 1 x Adapter /Standard SABS compliant Plug"]
["4.6"," 1 x Carry bag, and a set of rechargeable lithium batteries  for operation of the BP machine and separate charger if not onboard charging (built-in on-board preferable)"]
["5","Conformity Compliance (Please attach certificate)"]
[5.1,"The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attach certification"]
["5.2","OEM must comply and be certified on ISO 13485  quality standards, attach proof of compliance"]
["5.2.1","Please provide unique ref number of the ISO 13485 ceriticate:"]
["5.3","Model quoted for must be EC certified or equivalent. Attach a copy of certification"]
["5.3.1","Please provide unique ref number of the EC ceriticate:"]
["5.4","The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2"]
["6","COMPREHENSIVE MAINTENANCE "]
["6.1","A 24 month warranty inclusive maintenance and service/Replacement (Year 1 to Year 2) should be quoted for, as per the SCC document"]
["6.1.1","THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included.\nThis 2-year warranty will commence after formal Commissioning and handover of the equipment."]
["7","General  "]
["7.1","Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed"]
["7.2","The unit on offer must be the latest technology. "]
[8,"Mandatory/replacement accessories (pricing to be made available for future procurement)"]
["8.1"," 1x  Reusable latex-free small adult cuff"]
["8.2"," 1 x Reusable  latex-free standard adult cuff"]
["8.3"," 1 xReusable  latex-free obese cuff"]
["8.4"," 1 x NIBP hose"]
["8.5"," 1 x Adapter "]
["8.6"," 1x Rechargeable lithium battery"]

{"sheet":"HB Meter","index":1,"n_rows":54,"n_cols":4,"rows_kept":54,"merged":[[4,2,4,3],[5,2,5,3],[6,2,6,3]]}
[null,"Haemoglobin Meter","Complies Yes/No","Provide your answers/comments in this Column.  You are advised to be straight to the point."]
[1,"DESCRIPTION"]
["1.1","This specification establishes the requirements, supply, delivery, end user training, demonstration, installation and commissioning of a portable Haemoglobin meter able to measure accurate results covering the whole measurement range of <4.5- >26 g/dl (0-26 g/dl or 0-16 mmol/L)"]
[2,"Product Offered:"]
["2.1","Make:"]
["2.2","Model:"]
["2.3","Initial Manufacturing year of Model:"]
[3,"SPECIFICATIONS"]
["3.1","The unit offered must be robust, lightweight, compact and user friendly."]
["3.2","The unit must be accurate. Shorlisted units will undergo a trial as part of evaluation."]
["3.3","Bider to provide the following regarding the unit:"]
["3.4","Number of test results that can be stored: Onboard memory: 1000 tests"]
["3.5","Operating conditions: 10 - 40 degrees celcius; relative humidity: 35 - <90%   "]
["3.6"," Storage conditions: 0 - 50 degrees celcius, relative humidity: < 90%"]
["3.7","Minimum sample volume 10 - 20 microlitres"]
["3.8","Weight (in grams)"]
["3.9","The unit must be able to analyze capillary, venous or arterial whole blood"]
["3.10","Time to obtain results: < 30 seconds)"]
["3.12","The unit must be battery/electrically operated (re-chargeable or Alkaline batteries) "]
["3.13","The unit must have a built in memory to record the measurements."]
["3.14","The unit must be supplied with a carry case for protection."]
["3.15","The unit must be operated without any additional eye protection for both the user and patient"]
["3.16","The meter must have a digital display which is visible under all lighting conditions"]
["3.17","The unit must be able to display haemoglobin test in g/l, g/dl and mmol/l"]
["3.18","The above units must be user selectable as per their preference"]
["3.17","With measurement range of < 4.5 - >26 g/dl or/and 0 - 16 mmol/L "]
["3.18","The Unit should be provided complete with all accessories  for operation."]
["3.19","The unit must be Factory calibrated and and have built-in auto self calibration"]
["3.20","Automatic shutdown if not in use: < 6 minutes"]
[4,"Standard Accessories and Conusmables"]
["4.1","Carry case/bag for safe-keeping"]
["4.2","Battery (alkaline/lithium rechargeable batteries)"]
["4.3","Test strips (Pk of 50)"]
["4.4","Bluetooth, USB and/or other data transfer options"]
["4.5","Single-use disposable lancets "]
[]
["5","Conformity Compliance (Please attach certificate)"]
["5.1","The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification"]
["5.2","OEM must comply and certified on ISO 13485  quality standards, attach proof of compliance"]
["5.2.1","Please provide unique ref number of the ISO 13485 ceriticate:"]
["5.3","Model quoted for must be EC certified. Attach a copy of certification"]
["5.3.1","Please provide unique ref number of the EC ceriticate:"]
["6","COMPREHENSIVE MAINTENANCE "]
["6.1","A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price"]
["6.1.1","THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/on-board Calibrations, Software updates and upgrades to be included.\nThis 2-year warranty will commence after formal Commissioning and handover of the equipment."]
["7","General  "]
["7.1","Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed"]
["7.2","The unit on offer must be the latest technology. "]
["8","MANDATORY Accessories and Consumables (pricing to be made available for future procurement)"]
["8.1","Battery (alkaline/lithium rechargeable batteries)"]
["8.2","Test Strips"]
["8.3","quality control Solutions"]
["8.4","Bidder to list any other Consumables or Accessories (Not listed above) for the model on offer below"]
["8.5","Single-use disposable lancets"]

{"sheet":"Multiparameter","index":2,"n_rows":75,"n_cols":4,"rows_kept":75,"merged":[[4,2,4,3],[5,2,5,3],[6,2,6,3]]}
[null,"Multi-parameter Machine","Complies Yes/No","Provide your answers in this Column.  You are advised to be straight to the point."]
[1,"DESCRIPTION"]
[1.1,"Multi parameter testing machine with the following tests: ketones, cholesterol, lactate, uric acid "]
[2,"Product Offered:"]
[2.1,"Make:"]
[2.2,"Model:"]
["2,3","Initial Manufacturing year of Model:"]
["3","SPECIFICATIONS"]
["3.1","The unit offered must be robust, lightweight, compact and user friendly."]
["3.2","The unit must be accurate. Shorlisted units will undergo a trial as part of evaluation."]
["4","Bider to provide the following regarding the unit:"]
["4.1","Memory Capacity: minimum 100 patients (Data to be captured and uploaded)"]
["4.2","Approximate Dimensions:  Bidders to supply information"]
["4.3","Unit Must have Bluetooth or other Wireless equivalent"]
["4.4","Unit must have USB/micro"]
["4.5","Operating Conditions: 8°C~45°C, relative humidity:10%-90% "]
["4.6","Storage Conditions: -20 ºC ~ +60 ºC (Meter); 2ºC ~ 30ºC (Strips)"]
["4.7","Approcimate weight: Bidders to supply information"]
["4.8","The unit must be battery operated (rechargeable lithium/alkaline battery)."]
["4.9","Accuracy: A minimum of 1 local correlation trials minimum)(Proof to be supplied for model of machine) Specificity and sensitivity should correlate with laboratory-based results. Bidder must be SAHPRA licensed for medical device"]
["4.10","The unit must be supplied with a carry case for protection."]
["4.11","The meter must have a digital display which is visible under all lighting conditions"]
["5","KETONE "]
["5.1","Enzyme type:HBD"]
["5.2","Sample size: 0.8 µL"]
["5.3","Reaction Time: 10 seconds"]
["5.4","Measurement Range: 0.1 ~ 8.0 mmol/L"]
["5.5","Hematocrit Range: 10% - 70%"]
["5.6","Precision:  1 mmol/L, SD < 0.1 mM; > 1 mmol/L, CV < 7.5%"]
["5.7","Package: Vial pack 50"]
["6","TOTAL CHOLESTEROL  "]
["6.1","Sample size: 3.0 µL"]
["6.2","Reaction Time: 60 seconds"]
["6.3","Measurement Range: 100 ~ 400 mg/dL"]
["6.4","Hematocrit Range: 20% - 60%"]
["6.5","Precision: CV < 7.5%"]
["6.6","Package: Vial Pack 50"]
["7","URIC ACID  "]
["7.1","Sample size:0.5 µL"]
["7.2","Reaction Time: 15 seconds"]
["7.3","Measurement Range:3-20 mg/dL (178-1190 µM)"]
["7.4"]
["7.5","Hematocrit Range: 20% - 60%"]
["7.6","Precision: 5 mg/dL, SD<0.5 mg/dL; >5 mg/dL, CV<7.5% (Ù297 µM, SD<30 µM; >297 µM, CV<7.5%)(Ù0.29 mM, SD<0.03 mM; >0.29 mM, CV<7.5%)"]
["7.7","Package: Vial pack 50"]
["8","LACTATE  "]
["8.1","Sample size: 0.8 µL"]
["8.2","Reaction Time: 5 seconds"]
["8.3","Measurement Range: 0.3 - 22 mmol/L"]
["8.4","Hematocrit Range: 10% - 65%"]
["8.5","Precision: Ù3mmol/L, SD<0.3mM; >3mmol/L, CV<7.5%"]
["8.6","Package: Vial pack 50"]
["9","Standard Accessories and Conusmables"]
["9.1","Battery   "]
["9.2","Test Strips per test parameter (Pk of 50)"]
["9.3","The unit on offer must be the latest technology. "]
["10","Conformity Compliance (Please attach certificate)"]
["10.1","The unit must comply with an acceptable international electrical safety standard such as IEC 61010-1 for medical equipment, attached certification","Complies Yes/No","Provide your answers in this Column. You are advised to be straight to the point"]
["10.2","Original Equipment Manufacturer (OEM) must comply and certified on ISO 13485  quality standards, attach proof of compliance"]
["10.3","Please provide unique ref number of the ISO 13485 ceriticate:"]
["10.4","Model quoted for must be EC certified. Attach a copy of certification"]
["10.5","Please provide unique ref number of the EC ceriticate:"]
["10.6","OEM must comply and certified on ISO 15197:2013  quality standards, attach proof of compliance"]
["10.7","Please provide unique ref number of the ISO 15197:2013 ceriticate:"]
["11","COMPREHENSIVE MAINTENANCE "]
["11.1","A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price/ Replacement"]
["11.2","THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included.\nThis 2-year warranty will commence after formal Commissioning and handover of the equipment."]
["12","General  "]
["12.1","Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed"]
["12.2","The unit on offer must be the latest technology. "]
["13","MANDATORY Accessories and Consumables"]
["13.1","Battery   "]
["13.2","Test Strips"]
["13.3","quality control Solutions"]
["13.4","Bidder to list any other Consumables or Accessories or Software (Not listed above) for the model on offer below"]

{"sheet":"Dual GLu &HbA1c","index":3,"n_rows":65,"n_cols":8,"rows_kept":65,"merged":[[4,2,4,3],[5,2,5,3],[6,2,6,3]]}
[null,"Dual Glucose HBA1c Meter","Complies Yes/No","Provide your answers in this Column.  You are advised to be straight to the point."]
[1,"DESCRIPTION "]
[1.1,"Portable dual handheld analyser with the following tests: Glucose and HBA1C"]
[2,"Product Offered:"]
[2.1,"Make:"]
[2.2,"Model:"]
["2,3","Initial Manufacturing year of Model:"]
["3","SPECIFICATIONS"]
["3.1","The unit offered must be robust, lightweight, compact and user friendly. "]
["3.2"," Shorlisted units will undergo a trial as part of verification"]
["3.3","Bidder to provide the following regarding the unit:"]
["3.4","Memory Capacity: minimum100 patients"]
["3.5","Approximate Dimensions: Companies to supply"]
["3.6","unit must have Bluetooth or other Wireless equivalent"]
["3.7","Unit must have USB/micro",{"$nulls":5}," "]
["3.8","Operating conditions: Glucose: 5-40°C, 10%-90% R.H\nHbA1c: 10-40°C, 30%-75% R.H"]
["3.9","Storage conditions : -20 ºC ~ +55 ºC (Meter); 2ºC ~ 30ºC (Strips)  HbA1c: 4ºC ~ 30ºC"]
["3.10","Approximate weight: Company to supply"]
["3.11","The unit must be battery operated. Batteries must be Lithium based Rechargable /alkaline"]
["3.12","Accuracy: A minimum of 1 local correlation trials minimum)(Proof to be supplied for model of machine)Specificity and sensitivity should correlate with laboratory-based results. Bidder must be SAHPRA licensed for medical device"]
["3.13","The unit must be supplied with a carry case for protection."]
["3.14","The meter must have a digital display which is visible under all lighting conditions"]
["4","Glucose "]
["4.1","Sample size: 0.8 µL"]
["4.2","Reaction Time: max 10 seconds"]
["4.3","Measurement Range: 0.6~33.3 mmol/L"]
["4.4","Precision:  ≥5.55 mmol/L, biasis<+15%"]
["4.5","Package: Vial pack 50"]
["5","HBA1C "]
["5.1","Sample size: 3.0 µL"]
["5.2","Reaction Time: max 7 minutes "]
["5.3","Measurement Range: 4.0%-14.0%\n20.2-129.5 mmol/mol\neAG: 67.9-354 mg/dL"]
["5.4","Precision: CV < 5%"]
["5.5","Package: Vial Pack 50"]
["5.6","Standard Accessories and Conusmables"]
["5.7","Battery   "]
["5.8","Test Strips per test parameter (Pk of 50)"]
["5.9","The unit on offer must be the latest technology. "]
["5.11","Conformity Compliance (Please attach certificate)"]
["5.12","The unit must comply with an acceptable international electrical safety standard such as IEC 61010-1 for medical equipment, attached certification","Complies Yes/No"]
["5.13","OEM must comply and certified on ISO 13485  quality standards, attach proof of compliance"]
["5.14","Please provide unique ref number of the ISO 13485 ceriticate:"]
["5.15","Model quoted for must be EC certified. Attach a copy of certification"]
["5.16","Please provide unique ref number of the EC ceriticate:"]
["5.17","COMPREHENSIVE MAINTENANCE "]
["5.18","A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price/ Replacement"]
["5.19","THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included.\nThis 2-year warranty will commence after formal Commissioning and handover of the equipment."]
["5.20","General  "]
["5.21","Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed"]
["5.22","The unit on offer must be the latest technology. "]
["5.23","Mandatory Accessories and Consumables"]
["5.24","Battery   "]
["5.25","Test Strips"]
["5.26","quality control Solutions"]
["5.27","Bidder to list any other Consumables or Accessories (Not listed above) "]
["6","Software :"]
["6.1","Bidders to supply dowloadable/installable software on smart mobile/computer devices as the endusers may require. This should be compatible with  bluetooth/USB or other equivalent means of data transfer/capture. Shortlisted Bidders will be required to demonstrate the software interface to the Bid Evaluation Commitee"]
["6.2.","Key features of the software should at minimum include:"]
["6.2.1","Type of Test(Glucose,HbA1C ...etc)"]
["6.2.2","Patient Name or Identifier"]
["6.2.3","Health Care worker Name or Identifier"]
["6.2.4","Permissions for different Health Care Workers/Doctor"]
["6.2.5","Locations( Clinic/district/province... etc)"]
["6.2.6","Short Report/intervention/s"]
["6.2.7","Date and Time"]

{"sheet":"HbA1c","index":4,"n_rows":49,"n_cols":8,"rows_kept":49,"merged":[[4,2,4,3],[5,2,5,3],[6,2,6,3]]}
[null,"HbA1c Meter ","Complies Yes/No","Provide your answers in this Column.  You are advised to be straight to the point."]
[1,"DESCRIPTION "]
[1.1,"Portable /handheld HbA1c test analyser "]
[2,"Product Offered:"]
[2.1,"Make:"]
[2.2,"Model:"]
["2,3","Initial Manufacturing year of Model:"]
["3","SPECIFICATIONS"]
["3.1","The unit offered must be robust, lightweight, compact and user friendly."]
["3.2"," Shorlisted units will undergo a trial as part of verification"]
["3.3","Bidder to provide the following regarding the unit:"]
["3.4","Memory Capacity: minimum100 patients"]
["3.5","Approximate Dimensions: Bidders to supply"]
["3.6","unit must have Bluetooth or other wireless equivalent"]
["3.7","Unit must have USB/micro",{"$nulls":5}," "]
["3.8","Operating conditions: 10-40°C, 30%-75% R.H"]
["3.9","Storage conditions : 4ºC ~ 30ºC(Room temperature)"]
["3.10","Approximate weight: bidders to supply"]
["3.11","The unit must be battery operated. (Rechargable /alkaline)"]
["3.12","Accuracy: A minimum of 1 local correlation trials minimum)(Proof to be supplied for model of machine)Specificity and sensitivity should correlate with laboratory-based results. Bidders must be SAHPRA licenced for supply of medical devices."]
["3.13","The unit must be supplied with a carry case for protection."]
["3.14","The meter must have a digital display which is visible under all lighting conditions"]
["4","HBA1C "]
["4.1","Sample size: 3.0 µL"]
["4.2","Reaction Time: max 7 minutes "]
["4.3","Measurement Range: 4.0%-15.0%\n20-140 mmol/mol\neAG: 67.9-354 mg/dL"]
["4.4","Package\" strips package of 50"]
["4.5","Test Strips to be compatible with dual meter"]
["5","Standard Accessories and Conusmables"]
["5.1","Battery   "]
["5.2","Test Strips (Pk of 50)"]
["5.3","The unit on offer must be the latest technology. "]
["6","Conformity Compliance (Please attach certificate)"]
["6.1","The unit must comply with an acceptable international electrical safety standard such as IEC 61010-1 for medical equipment, attached certification"]
["6.2","OEM must comply and certified on ISO 13485  quality standards, attach proof of compliance"]
["6.3","Please provide unique ref number of the ISO 13485 ceriticate:"]
["6.4","Model quoted for must be EC certified. Attach a copy of certification"]
["6.5","Please provide unique ref number of the EC ceriticate:"]
[7,"COMPREHENSIVE MAINTENANCE "]
["7.1","A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price/ Replacement"]
["7.2","THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included.\nThis 2-year warranty will commence after formal Commissioning and handover of the equipment."]
[8,"General  "]
["8.1","Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed"]
["8.2","The unit on offer must be the latest technology. "]
[9,"Mandatory Accessories and Consumables"]
["9.1","Battery "]
["9.2","Test Strips"]
["9.3","quality control Solutions"]
["9.4","Bidder to list any other Consumables or Accessories (Not listed above) "]

{"sheet":"Glucometer","index":5,"n_rows":56,"n_cols":8,"rows_kept":56,"merged":[[4,2,4,3],[5,2,5,3],[6,2,6,3]]}
[null,"Glucose meter","Complies Yes/No","Provide your answers in this Column.  You are advised to be straight to the point."]
[1,"DESCRIPTION "]
[1.1,"Portable Glucose testing device"]
[2,"Product Offered:"]
[2.1,"Make:"]
[2.2,"Model:"]
["2,3","Initial Manufacturing year of Model:"]
["3","SPECIFICATIONS"]
["3.1","The unit offered must be robust, lightweight, compact and user friendly."]
["3.2"," Shorlisted units will undergo a trial as part of verification"]
["3.3","Bidder to provide the following regarding the unit:"]
["3.4","Memory Capacity: minimum100 patients"]
["3.5","Approximate Dimensions: Companies to supply"]
["3.6","unit must have Bluetooth "]
["3.7","Unit must have USB/micro",{"$nulls":5}," "]
["3.8","Operating conditions: Glucose: 5-45°C, 10%-90% R.H"]
["3.9","Storage conditions : -20 ºC ~ +55 ºC (Meter); 2ºC ~ 32ºC (Strips)  "]
["3.10","Approximate weight: Bidders to supply"]
["3.11","The unit must be battery operated. Battery must be Rechargable "]
["3.12","Accuracy: A minimum of 1 local correlation trials minimum)(Proof to be supplied for model of machine)Specificity and sensitivity should correlate with laboratory-based results. Bidders must be SAHPRA licensed."]
["3.13","The unit must be supplied with a carry case for protection."]
["3.14","The meter must have a digital display which is visible under all lighting conditions"]
["4","Glucose "]
["4.1","Sample size: 0.8 µL"]
["4.2","Reaction Time: max 10 seconds"]
["4.3","Measurement Range: 0.6~33.3 mmol/L"]
["4.4","Package: Strips pack of 50"]
["4.5","Strips to be compatible with dual meter"]
["5.","Standard Accessories and Conusmables"]
["5.1","Battery   "]
["5.2","Test Strips (Pk of 50)"]
["5.3","The unit on offer must be the latest technology. "]
["6","Conformity Compliance (Please attach certificate)"]
["6.1","The unit must comply with an acceptable international electrical safety standard such as IEC 61010-1 for medical equipment, attach certification"]
["6.2","OEM must comply and certified on ISO 13485  quality standards, attach proof of compliance"]
["6.3","Please provide unique ref number of the ISO 13485 ceriticate:"]
["6.4","Model quoted for must be EC certified. Attach a copy of certification"]
["6.5","Please provide unique ref number of the EC ceriticate:"]
[7,"COMPREHENSIVE MAINTENANCE "]
["7.1","A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price/ Replacement"]
["7.2","THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included.\nThis 2-year warranty will commence after formal Commissioning and handover of the equipment."]
[8,"General  "]
["8.1","Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed"]
["8.2","The unit on offer must be the latest technology. "]
[9,"Mandatory Accessories and Consumables"]
["9.1","Battery   "]
["9.2","Test Strips"]
["9.3","quality control Solutions"]
["9.4","Bidder to list any other Consumables or Accessories (Not listed above) "]
[10]
["10.1"]
["10.2"]
["10.3"]
["10.4"]
["10.5"]
["10.6"]

{"sheet":"Cholesterol","index":6,"n_rows":52,"n_cols":4,"rows_kept":52,"merged":[[4,2,4,3],[5,2,5,3],[6,2,6,3]]}
[null,"Cholesterol Meter","Complies Yes/No","Provide your answers in this Column.  You are advised to be straight to the point."]
[1,"DESCRIPTION"]
[1.1,"Testing machine For Cholesterol "]
[2,"Product Offered:"]
[2.1,"Make:"]
[2.2,"Model:"]
["2,3","Initial Manufacturing year of Model:"]
["3","SPECIFICATIONS"]
["3.1","The unit offered must be robust, lightweight, compact and user friendly. "]
["3.2","The unit must be accurate. Shorlisted units will undergo a trial as part of evaluation."]
["4","Bider to provide the following regarding the unit:"]
["4.1","Memory Capacity: minimum 100 patients (Data to be captured and uploaded)"]
["4.2","Approximate Dimensions: Bidders to supply"]
["4.3","Unit Must have Bluetooth /other equivalent wireless connectivity"]
["4.4","Unit must USB/micro"]
["4.5","Operating Contitions: 8°C-45°C, 10%-90% R.H"]
["4.6","Storage Conditions: -20 ºC ~ +60 ºC (Meter); 2ºC ~ 30ºC (Strips)"]
["4.7","Approximate weight: Bidders to supply information"]
["4.8","The unit must be rechargeable lithium/alkaline battery operated."]
["4.9","Accuracy: A minimum of 1 local correlation trials minimum)(Proof to be supplied for model of machine) Specificity and sensitivity should correlate with laboratory-based results. Bidder must be SAHPRA licensed for medical device"]
["4.10","The unit must be supplied with a carry case for protection."]
["4.11","The meter must have a digital display which is visible under all lighting conditions"]
["5","TOTAL CHOLESTEROL "]
["5.1","Sample size: 3.0 µL"]
["5.2","Reaction Time: 60 seconds"]
["5.3","Measurement Range: 100 ~ 400 mg/dL"]
["5.4","Precision: CV < 7.5%"]
["5.5","Package: Vial Pack 50"]
["5.6","Test strips to be compatible with multi-parameter machine"]
["6","Standard Accessories and Conusmables"]
["6.1","Lithium /Rechargeable Battery  /alkaline batteries"]
["6.2","Test Strips per test parameter (Pk of 50)"]
["6.3","The unit on offer must be the latest technology. "]
["7","Conformity Compliance (Please attach certificate)"]
["7.1","The unit must comply with an acceptable international electrical safety standard such as IEC 61010-1 for medical equipment, attached certification","Complies Yes/No","Provide your answers in this Column. You are advised to be straight to the point"]
["7.2","OEM must comply and certified on ISO 13485  quality standards, attach proof of compliance"]
["7.3","Please provide unique ref number of the ISO 13485 ceriticate:"]
["7.4","Model quoted for must be EC certified. Attach a copy of certification"]
["7.5","Please provide unique ref number of the EC ceriticate:"]
["7.6","OEM must comply and certified on ISO 15197:2013  quality standards, attach proof of compliance"]
["7.7","Please provide unique ref number of the ISO 15197:2013 ceriticate:"]
["8","COMPREHENSIVE MAINTENANCE "]
["8.1","A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price/ Replacement"]
["8.2","THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included.\nThis 2-year warranty will commence after formal Commissioning and handover of the equipment."]
["9","General  "]
["9.1","Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed"]
["9.2","The unit on offer must be the latest technology. "]
["10","Mandatory accessories and consumables"]
["10.2","Rechargable battery"]
["10.2","Test Strips"]
["10.3","Quality control Solutions"]
["10.4","Bidder to list any other Consumables or Accessories (Not listed above) for the model on offer "]

{"sheet":"Ketones","index":7,"n_rows":53,"n_cols":4,"rows_kept":53,"merged":[[4,2,4,3],[5,2,5,3],[6,2,6,3]]}
[null,"Ketones Meter","Complies Yes/No","Provide your answers in this Column.  You are advised to be straight to the point."]
[1,"DESCRIPTION"]
[1.1,"Testing machine  for Ketones "]
[2,"Product Offered:"]
[2.1,"Make:"]
[2.2,"Model:"]
["2,3","Initial Manufacturing year of Model:"]
["3","SPECIFICATIONS"]
["3.1","The unit offered must be robust, lightweight, compact and user friendly. "]
["3.2","The unit must be accurate. Shorlisted units will undergo a trial as part of evaluation."]
["4","Bider to provide the following regarding the unit:"]
["4.1","Memory Capacity: minimum 100 patients (Data to be captured and uploaded)"]
["4.2","Approximate Dimensions: Bidders to supply "]
["4.3","Unit Must have Bluetooth or Equivalent wireless connectivity"]
["4.4","Unit must USB/micro"]
["4.5","Operating Contitions: 8°C~45°C, 10%-90% R.H"]
["4.6","Storage Conditions: -20 ºC ~ +60 ºC (Meter); 2ºC ~ 32ºC (Strips)"]
["4.7","Approximate weight: Biddes to supply "]
["4.8","The unit must be battery operated. (rechargeable lithium/alkaline?)"]
["4.9","Accuracy: A minimum of 1 local correlation trials minimum)(Proof to be supplied for model of machine) Specificity and sensitivity should correlate with laboratory-based results.  Bidder must be SAHPRA licensed for medical device"]
["4.10","The unit must be supplied with a carry case for protection."]
["4.11","The meter must have a digital display which is visible under all lighting conditions"]
["5","KETONE "]
["5.1","Enzyme type:HBD"]
["5.2","Sample size: 0.8 µL"]
["5.3","Reaction Time: 10 seconds"]
["5.4","Measurement Range: 0.1 ~ 8.0 mmol/L"]
["5.6","Precision:  1 mmol/L, SD < 0.1 mM; > 1 mmol/L, CV < 7.5%"]
["5.7","Package: strips pack 50"]
["5.8","Test strips to be compatible with multi-parameter machine"]
["6","Standard Accessories and Conusmables"]
["6.1","Rechargeable /alkaline Battery    "]
["6.2","Test Strips per test parameter (Pk of 50)"]
["6.3","The unit on offer must be the latest technology. "]
["7","Conformity Compliance (Please attach certificate)"]
["7.1","The unit must comply with an acceptable international electrical safety standard such as IEC 61010-1 for medical equipment, attached certification","Complies Yes/No","Provide your answers in this Column. You are advised to be straight to the point"]
["7.2","OEM must comply and certified on ISO 13485  quality standards, attach proof of compliance"]
["7.3","Please provide unique ref number of the ISO 13485 ceriticate:"]
["7.4","Model quoted for must be EC certified. Attach a copy of certification"]
["7.5","Please provide unique ref number of the EC ceriticate:"]
["7.6","OEM must comply and certified on ISO 15197:2013  quality standards, attach proof of compliance"]
["7.7","Please provide unique ref number of the ISO 15197:2013 ceriticate:"]
["8","COMPREHENSIVE MAINTENANCE "]
["8.1","A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price/ Replacement"]
["8.2","THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included.\nThis 2-year warranty will commence after formal Commissioning and handover of the equipment."]
["9","General  "]
["9.1","Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed"]
["9.2","The unit on offer must be the latest technology. "]
["10","Mandatory Accessories and Consumables"]
["10.1","Rechargeable /alkaline Battery  "]
["10.2","Test Strips"]
["10.3","quality control Solutions"]
["10.4","Bidder to list any other Consumables or Accessories (Not listed above) for the model on offer below"]