Specifications for microplate washer:

- Must be able to accept 96 well plates.

- Must be supplied with one wash buffer bottle (volume not less than 1L) and one rinse bottle (volume not less than 1L).

- Must be supplied with an 8way wash head.

- Residual volume must not exceed 2µL.

- Must be supplied with an aerosol cover.

- Must have an USB port for data transfer.

- Quotation must include all on-board software necessary for operation.

- Must have at least 90 wash protocols.

- Dimensions: must not exceed 35cm (H) x 39cm (W) x 25cm (D).

- Dispensing volume range must cover at least 100-400 µL

- Dispensing volume error must not exceed 6% for 300µL dispensing volumes when used with an 8way washing head.

- Dispensing precision must be 4% or smaller when used with an 8way washing head and 300µL washing volumes.

- Must be compatible with 240V, 50Hz, single phase electricity.

- Instrument must have a minimum warranty of 12 months.

- Quotation must include delivery, commissioning and training of 2 researchers or laboratory technicians.

- Spares must be available for at least 10 years from date of commissioning. Service provider must have a workshop and factory trained technicians. Training certificates of technicians must be submitted with quotations and customer must be allowed a site visit to verify that service provider has a workshop.

- Washer must be supplied with fitted surge protection plug.

|  | Requirement | Comply |  |
| --- | --- | --- | --- |
|  | Specifications | Yes | No |
| 1 | Must be able to accept 96 well plates. |  |  |
| 2 | Must be supplied with one wash buffer bottle (volume not less than 1L) and one rinse bottle (volume not less than 1L). |  |  |
| 3 | Must be supplied with an 8way wash head. |  |  |
| 4 | Residual volume must not exceed 2µL. |  |  |
| 5 | Must be supplied with an aerosol cover. |  |  |
| 6 | Must be supplied with an aerosol cover. |  |  |
| 7 | Quotation must include all on-board software necessary for operation. |  |  |
| 8 | Must have at least 90 wash protocols. |  |  |
| 9 | Dimensions: must not exceed 35cm(H) x 39cm(W) x 25cm(D). |  |  |
| 10 | Dispensing volume range must cover at least 100-400 µL |  |  |
| 11 | Dispensing volume error must not exceed 6% for 300µL dispensing volumes when used with an 8way washing head. |  |  |
| 12 | Dispensing precision must be 4% or smaller when used with an 8way washing head and 300µL washing volumes. |  |  |
| 13 | Must be compatible with 240V, 50Hz, single phase electricity. |  |  |
| 14 | Instrument must have a minimum warranty of 12 months. |  |  |
| 15 | Quotation must include delivery, commissioning and training of 2 researchers or laboratory technicians. |  |  |
| 16 | Spares must be available for at least 10 years from date of commissioning. Service provider must have a workshop and factory trained technicians. Training certificates of technicians must be submitted with quotations and customer must be allowed a site visit to verify that service provider has a workshop. |  |  |
| 17 | Washer must be supplied with fitted surge protection plug. |  |  |