LABORATORY EQUIPMENT SPECIFICATIONS

- 1 x CONCRETE TESING MACHINE – AUTOMATIC

| Capacity | 3000 kN |
| --- | --- |
| Standard | EN |
| The roughness value for texture of loading and auxiliary platens | ≤ 3.2 µm |
| Lower platens dimensions (dia.) | 300 mm |
| Upper platens dimensions (dia.) | 300 mm |
| Maximum vertical clearance between platens | 340 mm |
| Piston diameter | 300 mm |
| Maximum piston movement | 50 mm |
| Horizontal clearance | 425 mm |
| Maximum working pressure | 410 Bar |

- Automatic test cycle with standard rapid approach

- Graphical – LCD data acquisition and control system

- Menu driven software for easy operation

- Load vs time plot and instantaneous load rate displayed

- Accurate load rate control within ±5 from 1 kN/sec to 20 kN /sec

- Class 1 starting from % 10 of the full range

- Ready stress calculation for the standards of EN 12390-3, EN 12390-5, EN 12390-6,  EN 1338, EN 1340, EN 196

- SI, Metric units

- Data storage up to 256 tests

- RS 232 output for printer or PC (OR ANY INCLUDING USB AND HDMI)

- PC software for test control and advanced report printout

- 1 x FLEXURE TESTING MACHINE – DIGITAL

- Automatic test cycle with standard Rapid Approach

- Graphical – LCD Data Acquisition and Control System

- Menu Driven Software for easy operation

- Load vs time plot and instantaneous load rate displayed

- Accurate load rate control within ±5 from 1 kN/sec to 20 kN /sec

- Class 1 starting from % 10 of the full range

- Ready Stress Calculation for the standards of EN 12390-5 , EN 12390-6,                 EN 1338 ,EN 1340, ASTM C78, C293, C496

- SI, Metric Units

- Data storage up to 256 tests

- 1 x CONCRETE CUBE MOULDS

- Faces are machined flat to +/-0.2mm accuracy and finished to within 0.2mm.

- Cube moulds to be supplied complete with base plates

- Cube Mould (Cast Iron – for Mortar) 50mm

- Single Gang Cube Mould (Cast Iron – for Mortar) 50mm

- Three Gang Cube Mould (Mild Steel – for Cement) 70.6mm

- Cube Mould (Cast Iron – for Concrete) 75mm

- Single Gang Cube Mould (Cast Iron – for Concrete) 100mm

- Light Weight Cube Mould (Cast Iron – for Concrete) 100mm

- Clamp Type (8.5 KG) Four-Part Cube Mould (Cast Iron – for Concrete) 150mm

- Light Weight 8KG Cube Mould (Cast Iron – for Concrete) 150mm

- ISI Marked Cube Mould (Cast Iron -for Concrete) 150mm

- Clamp Type (18 KG) Four-Part

- Cube Mould (Cast Iron – for Concrete) 150mm Clamp Type (16 KG) Two-Part

- 1 x CYLINDRICAL MOULDS

The cylindrical mould is split vertically into two parts. The mean internal diameter is within +/-0.2mm and height is within +/-1mm. The ends are machined to +/-0.05 mm. The base plate and top plate are machined flat to +/-0.03mm.

- Mould cylinder, cast iron, 150mm dia x 150mm height.

- Mould cylinder, cast iron, 150mm dia x 300mm height.

- Mould cylinder, cast iron, 160mm dia x 320mm height.

- Mould cylinder, cast iron, 100mm dia x 200mm height.

- Mould cylinder, cast iron, 300mm dia x 600mm height.

- 1 x BEAM MOULDS

Made of Cast Iron / Mild Steel. Beam moulds are made of 4 plates assembled together. Each bean mould is supplied complete with base plate. Faces are machined flat to ± 0.2mm. And finished in size to 0.2mm.

- 100 mm x 100 mm x 500 mm

- 150 mm x 150 mm x 700 mm

- 150 mm x 150 mm x 750 mm

- 100 mm x 150 mm x 600 mm

- 150 mm x 150 mm x 600 mm

- 1 x SLUMP TEST APPARATUS

The slump cone has internal dimension 200 mm dia at base, 100 mm top dia and 300 mm height. Foot pieces can be fixed to the clamps on the base plate. The base plate has lifting handle for easy transportation.

One graduated steel rod 16 mm dia x 600 mm long rounded at both ends and graduated in mm.

Include scoop, trowel, brush, measuring tape or meter rule, and/or funnel.

- 1 x LABORATORY PAN MIXER

- All parts of the mixer are galvanized and painted with non-corrosive paint. The protection class of machine is IP55

- Dimension: 650x900x1200 mm (Dimensional volume of the mixing pan is 42 litres and the effective mixing capacity is 15 litres)

- The mixing drum should be tilted to empty the mixture without hassle and it can be removed for easy cleaning on completion of the mixing operation.

- Should equipped with rubber wheels which provides high mobility.

- 1 x LABORATORY DRUM TYPE MIXER

- Drum volume:  125 litres

- Yield:                75 litres of concrete

- All parts of the mixer are galvanized and painted with non-corrosive paint. The protection class of machine is IP55

- The mixing drum should be tilted to empty the mixture without hassle and it can be removed for easy cleaning on completion of the mixing operation.

- Should equipped with rubber wheels which provides high mobility.

- 1 x CONCRETE PERMEABILITY APPARATUS

- Concrete Permeability Apparatus 3 cell model for 100 mm cube.

- Concrete Permeability Apparatus 3 cell model for 100 mm cylinder.

- Concrete Permeability Apparatus 3 cell model for 150 mm cube.

- Concrete Permeability Apparatus 1 cell model for 150 mm cylinder.

- Concrete Permeability Apparatus 1 cell model for 300 mm cylinder

- 1 x ACCELERATED CURING TANK

- 24 Hour cycle from time of mixing.

- Controlled 35° C or 100° C ± 2° C Curing Temperature for concrete.

- Controlled 27° C ± 2° C Curing Temperature for grey cement.

- Curing tank to accommodate 150mm / 70.6mm cube moulds and up to 36/72   cube moulds can be cured in one go

- The interior is made from stainless steel

- The electronic programmer should control up to four test cycles with different thermal gradients and curing time to a defined temperature value for a complete automatic curing cycle.

- The control panel should include a 28 column thermal printer

The unit should come with the following moulds:

- Curing Tank for 6/12 moulds of 150 mm / 70.6 mm size

- Curing Tank for 12/24 moulds of 150 mm / 70.6 mm size

- Curing Tank for 24/48 moulds of 150 mm / 70.6 mm size

- Curing Tank for 36/72 moulds of 150 mm / 70.6 mm size

- 1 x DIGITAL WEIGHING BALANCE

- Single-block weighing technology to ensure greater precision

- Backlight display to provide clearer view of results

- Simple menu navigation with easy access to advanced features

- Overload protection to prevent damage due to misuse

- Weighing range [Max]: 30,000 g / 30 kg

- Readability [d]: 1 g

- Dimension - Dimension weighing plate [WxD]: 260 x 220 mm

- Electricity and/or Battery operation

- Calibritable

- 1 x TENSILE STRENGTH TESTER

- A loading Machine, double lever type, with steel scale marked from 
0-500 Newtons in 10 Newton division.

- Maximum loading capacity 5 kN.

- Automatic Loading system using Lead Shot. Lead shot 15 kg
supplied with the machine.

- Set of weights for weighing lead shot 
comprising one each for weighing

- One standard Briquette Mould with Base plate

- 1 x DIGITAL SIEVE SHAKER

- ¼ HP unit

- 1 x FIELD GPS

- Must be sunlight readable

- Barometric altimeter

- Large screen >2.6”

- Handling with gloves

- Waterproof outdoor GPS

- Wirelessly connects to smartphone

- Tough and rugged

- 2.6" sunlight-readable color screen

- High-sensitivity GPS and GLONASS  quad helix antenna  (Global Navigation Satellite System)

- 1-year BirdsEye Satellite Imagery subscription

- 3-axis compass with barometric altimeter

- plot barometric pressure over time

- Wireless connectivity via Bluetooth® technology¹ or ANT+®

- 3-axis tilt-compensated electronic compass

- the dual battery system

- Include TOPO Southern Africa PRO maps

- Include software to transform paper and electronic maps easily into downloadable maps

- Bluetooth connectable

- 1 x DYNAMIC CONE PENETROMETER

- Supplied with spanners, allen key, weight & cone.

- DCP 60° cone (Hardened)

- 8kg drop weight

- 575mm drop height

- 1meter bottom Rod

- 1.5 meter bottom rod

- 1m steel ruler

- 1.5meter steel ruler

- Steel careening case (no carry bag)

- 10m Measuring tape

- Manual in English

- 1 x ULTRASONIC THICKNESS GAUGE

- Measuring Range:	 1 -300MM

- Accuracy:		minimum 0.1mm

- Resolution		0.1mm

- Tube measuring:	30x3mm

- Auto calibration

- Corrosion thickness gaging

- Data storage

- IP 67

- Complete with all consumables & software (Batteries, Cables, Ultrasonic Couplant ect) to make this an working instrument unit on delivery.

- Calibration block for steel: 8 units (Step blocks) 1 - 8mm

- Measurements on painted and coated surfaces, Build in function to ignore coating thickness (Paint)

- User’s manual

- 1 x CONCRETE TEST HAMMER

- Measuring range: 10-60 MPa

- Impact energy: 2.207±0.1J（0.225Kgf.m)

- Length of spring stretch: 75 ±0.3mm

- The static friction of pointer slider: 0.65N~0.15N

- Radius of spherical tip: 25mm ±1mm

- The average rebound values on steel anvil: 80 ±2

- Housing dimensions: Φ54 * 280mm length

- Weight: ≈1Kg

- 1 x CRACK MEASURING MICROSCOPE

- High quality microscope designed for measuring crack widths in concrete members, masonry walls and other structures

- Should operate by an adjustable lamp unit and the image is focused by turning a knob

- The eyepiece scale can be turned through 360° to align with the direction of the crack or pitch under examination.

- The battery operated microscope to have at least 40x magnification and 
at least 4 mm measuring range with 0.02 mm subdivisions.

- 1 x COMPACTION FACTOR APPARATUS

- Consists of two conical hoppers attached on a cylinder, and everything is fixed on a robust steel frame

- Consists of two conical hoppers attached on a cylinder, and everything is fixed on a robust steel frame

- Each hopper is equipped with a hinge for a Quick Release mechanism

Additional Information:

The above specified goods should be delivered to:

Name of Institute     : ARC-IAE

Address                   : 141 Creswell Road

: Weavind Park

: Pretoria

CLOSING DATE : 23 November 2021@16h00

KINDLY ENSURE THAT YOUR COMPANY IS REGISTERED ON CENTRALISED SUPPLIER DATABASE IF NOT ,KINDLY REGISTER ON THIS WEBSITE www.csd.gov.za

The particulars of the guarantee that will apply to the goods quoted for, with particular regard to the period and extent of the warranty must be clearly stated.

Where services are required, service providers must submit documentation pertaining the relevant experience.

Your written quotation should be sent to:

Lungile Kubheka

E Mail: KubhekaL@arc.agric.za

012 842 4078

All price quotations that have a rand value of R30, 000-00 to R50 000 000.00 including VAT, will be evaluated by applying the 80/20 principle as prescribed by the Preferential Procurement Policy Framework Act 5 of 2000 and its Regulations. The lowest acceptable price will score 80 points, the 20 BBBEE points will be allocated as follows:

| B-BBEE Status Level of Contributor | 80/20 |
| --- | --- |
| 1 | 20 |
| 2 | 18 |
| 3 | 14 |
| 4 | 12 |
| 5 | 8 |
| 6 | 6 |
| 7 | 4 |
| 8 | 2 |
| Non-compliant Contributor | 0 |

Standard conditions:

- The validity of the quotations must be indicated.

- Prices quoted should be in South African Rand and inclusive of VAT costs such as delivery, insurance, taxes, etc.

- No price adjustments or amendment of the delivery particulars contained in paragraph 2 will be considered by the ARC.

- The supplier accepts full responsibility for the proper execution and fulfilment of the goods quoted for.

- ARC reserves the right to accept or reject any special terms and conditions that may qualify the goods/services to be provided.

- Quotes should be submitted on an official letterhead and duly signed

- Goods and services should be supplied upon receipt of a purchase order from the ARC

- The General Conditions of Contract issued by National Treasury are applicable.

- The ARC supply chain management code of conduct is applicable.

- SBD Forms must be signed and returned together with the quotation if your price is above R10 000.00, failure to comply will result to disqualification of your quotation.

- Your quotation must indicate the delivery date

Compulsory Requirements:

- Valid Tax Pin Number from SARS.

- CSD report showing a tax compliant status for the duration of the RFQ process.

- Complete and Sign All the SBD forms.

The ARC reserve the right to do due diligence on the quotations.

The ARC reserve the right to benchmark prices quoted.

Thank you in anticipation