1 | Page

[image: image1.png]



SPECIFICATIONS FOR SUPPLY, DELIVERY, INSTALLATION, DEMONSTRATION AND

COMMISSIONING OF A DIGITAL FLOOR STAND X-RAY MACHINE WITH HEIGHT ADJUSTABLE TABLE TOP FOR 
VICTORIA WEST CHC, NORTHERN CAPE DEPARTMENT OF HEALTH
COMPULSORY SITE VISIT
	
	SPECIFICATIONS
	COMPLY/DO NOT COMPLY

YES/NO
	BIDDERS COMMENTS

	
	TECHNICAL
	
	

	1.1  
	The bid price quoted must include the supply, installation, demonstration and commissioning of the equipment
	
	

	1.2
	State the make and model of the unit offered. The projected life span of the machine must be included
	
	

	1.3
	All companies must ensure that the machines offered are compatible with the hospital’s electrical and IT systems. 
	
	

	1.4 
	The unit must be suitable for general and trauma radiography.
	
	

	1.5
	Please include a list of all hospitals or clinics where the equipment is installed and include at least 3 reference letters 
	
	

	1.6
	The unit should be at an acceptable frequency type generator for workload at Victoria West CHC..
	
	

	1.7
	The machine must be able to store images for a period of at least 2 years.
	
	

	1.8
	The machine must be linked to the available PACS system. State DICOM conformity
	
	

	1.9
	Include anti -virus software for the duration of the warranty and maintenance period
	
	

	
	TECHNICAL SPECIFICATION

GENERATOR:
	
	

	2.1
	The unit should be a high frequency type generator.


	
	

	2.2
	The output must be at least 50kW or more. State maximum milliamps at different kilovolts
	
	

	2.3
	The unit must incorporate an electronic timer with a minimum exposure time. Please state
	
	

	2.4
	State if the unit includes automatic overload protection. Yes or No
	
	

	2.5
	Does the unit must include automatic mains compensation.
	
	

	2.6
	A milliamp meter or indicator must be provided
	
	

	2.7
	The kV, mA and time must be independently selectable and digitally displayed.
	
	

	2.8
	Bidders to state the range of the exposure timer. Acceptable range 0.001 sec – 6 sec (1ms – 6sec). Please state
	
	

	2.9
	Bidders to state kV and mA ranges for radiography. The kV range to be at least 40 to 150 kV and mA range at least 10 to 630mA. State KV selection
	
	

	2.10
	The unit must preferably be programmable. Supply details of the programming facility and the number of programs.
	
	

	2.11
	The offered unit must include automatic exposure control.
	
	

	
	
	
	

	
	ROTATING ANODE X‑RAY TUBE
	
	

	3.1
	The rotating anode tube must match the output of the generator. Please state
	
	

	
	
	
	

	
	BIDDERS TO SUPPLY THE FOLLOWING DETAILS
	
	

	4.1
	The maximum tube kV must be at least 150 kV

State the maximum tube mA
State the speed of anode rotation.

State the Anode heat storage capacity 
State the tube housing heat storage capacity
State the size of small focus 

State the size of large focus
State the inherent filtration o
Total filtration to be stated.

Anode angle must be large enough to provide a uniform beam over a field of 43cm x 43cm at a FFD of 100cm.
	
	

	
	
	
	

	
	GENERAL 
	
	

	5.1
	ELECTRICAL: One pair of high tension cables of suitable length must be supplied. The cable length must not restrict free movement of the tube for all radiographic procedures.
	
	

	
	COLLIMATOR
	
	

	6.1
	The unit must have a multi leaf collimator.
	
	

	
	
	
	

	
	TUBE
	
	

	7.1
	TUBE STAND: A floor mounted tube stand must be supplied that will provide a focal film distance of at least 120 cm to the Bucky table. The tube stand must not be integrated into the table. The longitudinal movement of the tube stand must extend beyond the table at both ends.
	
	

	7.2
	The unit must be fitted with electro‑magnetic brakes.
	
	

	7.3
	Bidders to indicate the focal film distance range. Markings indicating the focal film distance on the tube stand must be provided.      
	
	

	7.4
	A safety device for possible failure of the counter‑balance system must be incorporated.    


	
	

	7.5
	The unit must be suitable to function with an x‑ray Bucky table, the vertical Bucky and provide a facility for horizontal beam radiography.
	
	

	7.6
	Bidders to state:

i) Tube rotation around vertical axis

ii) Tube rotation around horizontal axis and

iii)            Range of vertical movement.

iv)            Range of longitudinal movement.
	
	

	7.7
	The unit must be suitable for radiography of patients on trolleys and therefore space must be available at both ends of the table – e.g. it must be possible to do a supine chest x-ray or a tibia and fibula x-ray on a hospital trolley.  State the longitudinal movement of the tube column.
	
	

	7.8
	Rotation of the tube to the lateral position for horizontal shoot through radiographs must be possible by means of a central electronic locking/unlocking device at the collimator. Units that have the locking/unlocking device on the tube stand close to the tube rail will not be accepted
	
	

	
	
	
	

	
	
	
	

	
	BUCKY TABLE
	
	

	8.1
	The Bucky table must have a flat, four way floating top.
	
	

	8.2
	The table must be electronically height controlled.  Foot switches must form part of the table.
	
	

	8.3
	The table must be electronically height controlled.  Foot switches must form part of the table.
	
	

	8.4
	The bidder to state dimensions of table top offered (length and width).
	
	

	8.5
	State the variable height of the table offered.
	
	

	8.6
	State the lateral and longitudinal movement of the table top.
	
	

	8.7
	The table offered must be able to carry patients with a weight of at least 150kg or more. 
	
	

	8.8
	State longitudinal movement of Bucky tray. Bidders must detail whether the distance relates to the outermost end of the Bucky tray or from center point to center point. A data sheet must be included to clarify this point.
	
	

	8.9
	The table‑top movement must be controlled with electro‑magnetic brakes. 
	
	

	8.10
	State the table‑top ‑ film distance       
	
	

	8.11
	State whether the Bucky assembly is be fitted with the oscillating secondary radiation grids suitable for direct radiography
	
	

	8.13
	Mattress must be supplied for table top
	
	

	
	
	
	

	
	VERTICAL BUCKY
	
	

	9.1
	State the make and model of the vertical Bucky offered.
	
	

	9.2
	The Bucky assembly must be fitted with a oscillating secondary radiation grid of 10: 1, and be suitable for chest in Bucky radiography.
	
	

	9.3
	Radiography of a patient who is in a seated position must be possible.
	
	

	9.4
	State if the Bucky assembly is fitted with the oscillating secondary radiation grids suitable for direct digital flat panel technology and the distance accommodated for chest radiography
	
	

	9.5
	The Bucky must be mounted centrally on the column with the choice of loading on either left or right
	
	

	9.6
	Vertical movement of the vertical Bucky must be via an electronic locking/unlocking switch. Analogue locking/unlocking levers or knobs will not be accepted
	
	

	9.7
	An overhang frame for out of Bucky radiography must be included that is compatible with erect Bucky
	
	

	9.8
	An overhang arm support must be included that is compatible with erect Bucky
	
	

	
	
	
	

	
	INSTALLATION
	
	

	10.1
	i)              De-installation and disposal of existing equipment (where applicable), including the removal to a place designated by the Hospital management must be included in the quote
ii)             Room preparation to accommodate the new equipment will be done separately after an official site visit and input of the end users. Cost to be included in the quote.
iii)            Delivery, installation and commissioning of equipment to be included in the quote for the x-ray machine.
	
	

	10.2
	Prior arrangements must be made with Health Technology Services with regard to de-installation and disposal of the old unit.
	
	

	10.3
	The proposed layout of the equipment must be discussed with the Department of Health prior any room alterations
	
	

	10.4
	The layout and the alterations must be shown on a plan that must be submitted with the quote.
	
	

	10.5
	The Bidders must satisfy themselves as to the environmental conditions. Any additional features that are required, such as air conditioning, must be detailed separately and included with the quote.
	
	

	10.6
	The bid price must include

 i)  The necessary QA test, any additional items, modifications or building alterations required to meet the requirements of the Directorate Radiation Control of the Department of Health.

ii) Must be final and inclusive of all the necessary components that make up the final bid price.
	
	

	
	UPGRADEABILITY:
	
	

	11.1
	All future upgrades (hardware and software) involving patient safety and removing software viruses from existing software must be supplied at no additional cost.
	
	

	11.2
	ANY UPGRADE BEFORE OR AFTER INSTALLATION OF THE EQUIPMENT INVOLVING ADDITIONAL COST MUST BE BROUGHT TO THE ATTENTION OF THE MANAGER, RADIOLOGY AND HEALTH TECHNOLOGY SERVICES.
	
	

	
	
	
	

	12
	DIGITAL DETECTOR STAND 
	
	

	12.1
	Removable detector.
	
	

	12.2
	The detector must at least be multipurpose detector with good ergonomics, combined with great flexibility. The detector must allow fast and efficient exposures as well as special angulations across the table. (QUOTE ON A WIRELESS AND WIRED DETECTOR).
	
	

	12.4
	State the active image size of the detector. It must ensure high projection flexibility even with large patients. State offer.
	
	

	12.5
	The detector must be able to be positioned horizontal or vertical.
	
	

	12.6
	The detector must also be able to be used for on table and bedside exposures. Please explain.
	
	

	12.7
	An electronic flat detector made of amorphous silicon  and cesium iodide scintillator ,for excellent image quality even

with low dose, is required.  State offer.
	
	

	12.8
	 Image matrix size must be stated. Please attach proof of compliance.
	
	

	12.9
	Provide the acquisition depth  conversion 
	
	

	12.10
	State the pixel size. Please attach proof of compliance


	
	

	12.11
	State the image resolution

Please attach proof of compliance
	
	

	12.12
	State the sensitivity range 
	
	

	12.13
	Inter changeable, oscillating grid for enhanced clinical application and superior image quality is requested. 

Stathe the grid ratio of the unit 
	
	

	12.14
	 Automatic exposure control of minimum of 3 chambers, to ensure correct dosage of any projection done, must be included.
	
	

	12.15
	A patient mattress must be included.
	
	

	12.16
	A lateral cassette holder for decubitus projections to hold cassettes ranging from 35 x 43cm.
	
	

	12.17
	A lateral support arm for lateral chest imaging AND  the stand to store it must be included.
	
	

	13
	MANUALS AND BROCHURES :
	
	

	13.1
	The successful bidder must include in their offer at no extra cost to the final bid price:

i)              Complete user Operation / Maintenance Manual …copies in English Language

ii)            Complete ORIGINAL Service / Repair Manual x 2 (two) Book / File; CD; DVD copies in English Language which MUST include the following information:

Fault Finding Guide, Circuit Diagrams / Schematics, Circuit Descriptions, and PCB Layouts, Calibration    Guide, Part Numbers and exploded diagram of Mechanical Parts / Panels.
	
	

	13.2
	The offer submitted must be supported by descriptive literature, colour pamphlets, colour brochures and technical data sheets applicable to the offer 

FAILURE TO SUBMIT THE ABOVE WILL RESULT IN THE BID NOT BEING CONSIDERED.
	
	

	
	
	
	

	14
	TRAINING IN THE CORRECT USE OF PRODUCTS:
	
	

	14.1
	The successful bidder must offer continuous training to staff (radiographers and medical physicists) in effective utilization of their products.  Wastage as a result of not effectively utilizing products must be immediately reported by the supplier to the Department of Health. When called for by the Department of Health Technology Services, the contractor must furnish the details sought after.
	
	

	14.2
	The successful bidder must provide the Health Technology Service’s in house Technicians, a demonstration of the product offered, full training in the calibration, maintenance, service and repair of the product down to PCB Level. N.B. The quality and level of the training must be equivalent to the manufacturer’s original factory training and any costs incurred to provide this training will be for the bidders account. A Certificate of Competency must be issued on completion of the training. The Training must be provided by the successful bidder to the Health Technology Services within three months from date of initial supply and delivery of the equipment to the end user.
	
	

	14.3
	The successful Bidder must at no extra cost provide additional ongoing training for end users and technical staff on the equipment offered.
	
	

	
	
	
	

	15
	RADIATION CONTROL LICENCE
	
	

	15.1
	Bidders must state the Radiation Control License number of the make and model of the equipment offered. If this type of equipment/apparatus appears on the schedule of Hazardous Substances, issued by the Directorate: Radiation Control of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid document. The license must be registered under the bidders name or the letter of Joint Venture must be submitted by the License holder where the license is not in the name of the bidder. 
	
	

	
	BIDDERS THAT NEGLECT TO SUBMIT A LICENCE WILL BE DISQUALIFIED.
	
	

	
	BIDDER TO STATE LICENCE NUMBER:_________________________
	
	

	
	
	
	

	16
	GUARANTEE PERIOD
	
	

	16.1
	Bidders must supply a twenty-four month guarantee against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including, where appropriate, Consumables), labor, travelling and accommodation.

The two-year guarantee must be included in the unit price of the equipment. Full breakdown of costing schedule must be given. 

The annual quality assurance and maintenance must be included
	
	

	17
	MAINTENANCE CONTRACT
	
	

	17.1
	Spare parts must be available for the specified life of the equipment, with a minimum of ten years.

Spare part kits should specify all spare parts that will be included and stored on-site or available in the country.

It must be guaranteed that no additional equipment, parts or software, excluding consumables, is required to operate the equipment specified in this tender. Specify
Please state predicted life span of the machine.
	
	

	17.2
	The 5 year maintenance plan must be costed and included for implementation after the 24 month warranty period is over. Please include the price schedule to determine the total cost of ownership. Monthly payments to start after the warranty period
	
	

	18
	SCOPE OF WORK TO MAKE ROOM GOOD
	
	

	18.1
	A scope of work to be agreed upon at the site visit. This will only entail work done to ensure floors, joinery and walls are in good state post installation. 
	
	


Specifications prepared and supplied by:

Ms. J. A. Du Plooy

DD: Radiography

…………………………………………
Specifications for Victoria West 2025/26
Specifications for Digital Floor stand at Victoria West Hospital 2025/26

