Incubator specifications

- Setting temperature range: 0 to +70 °C

- Working temperature range with light: +10 to +40 °C

- Working temperature range without light: from 0 (at least 20 below ambient temperature) to +70°C

- Setting accuracy temperature: 0.1 °C

- Temperature sensor 1 Pt100 sensor DIN class A in 4-wire-circuit

- adjustable electronic overtemperature monitor and mechanical temperature limiter

- ControlCOCKPIT SingleDISPLAY. Adaptive multifunctional digital PID-microprocessor controller with high-definition TFT-colour display

- Timer Digital backwards counter with target time setting, adjustable from 1 minute to 99 days

- Function SetpointWAIT the process time does not start until the set temperature is reached

- Calibration three freely selectable temperature values

- adjustable parameters temperature (Celsius), programme time, time zones, summertime/wintertime

- energy-saving heating-/cooling system integrated in the rear (heat pump principle)

- Door fully insulated stainless steel door with 2-point locking (compression door lock)

- Internals 2 stainless steel grid(s), electropolished

- Door inner glass door

- Volume 384 litres

- 2-year warranty

- Delivery, installation and calibration must be included