WATER LEVEL DATALOGGER SPECIFICATIONS

Date compiled: 18 August 2025

Solinst Levelogger 5 is the mostly recommended one

| WATER LEVEL DATALOGGER SPECIFICATIONS |  |
| --- | --- |
| Level Sensor: | Piezoresistive Silicon with Hastelloy Sensor |
| Accuracy: | ± 0.05% FS (Barologger 5: ± 0.05 kPa) |
| Stability of Readings: | Superior, low noise |
| Resolution: | 0.002% FS to 0.0006% FS |
| Units of Measure: | m, cm, ft., psi, kPa, bar, ºC. ºF (Barologger 5: psi, kPa, mbar, ºC, ºF) |
| Normalization: | Automatic Temperature Compensation |
| Temp. Comp. Range: | 0º to 50ºC (Barologger 5: -10 to +50ºC) |
| Temperature Sensor: | Platinum Resistance Temperature Detector (RTD) |
| Temp. Sensor Accuracy: | ± 0.05ºC |
| Temp. Sensor Resolution: | 0.003°C |
| Battery Life: | 10 Years - based on 1 reading/minute |
| Clock Accuracy: | ± 1 minute/year (-20ºC to 80ºC) |
| Operating Temperature: | -20ºC to 80ºC |
| Maximum # Readings: | 150,000 sets of readings |
| Memory Mode: | Slate and Continuous |
| Communication: | Optical high-speed: USB, SDI-12 57,600 bps with USB |
| Size: | 22 mm x 160 mm (7/8" x 6.3") |
| Weight: | 146 grams (5.2 oz) |
| Corrosion Resistance: | Baked-on coating using polymerization technology (inside and out) |
| Other Wetted Materials: | Delrin®, Viton®, 316L stainless steel, Hastelloy, PFAS-free PTFE coating |
| Sampling Modes: | Linear, Event & User-Selectable with Repeat Mode, Future Start, Future Stop, Real-Time View |
| Measurement Rates: | 1/8 sec to 99 hrs |
| Barometric Compensation: | Software Wizard and one Barologger 5 in local area (approx. 30 km/20 miles radius) |