[bookmark: _GoBack]INCUBATOR WITHOUT FAN 161L, IN160
Quantity – 3
Project number – P09000339
Specifiations:

Temperature
	Setting temperature range
	+20 to +80 °C

	Working temperature range
	at least 5 above ambient temperature to +80 °C

	Setting accuracy temperature
	0.1 °C

	Temperature sensor
	1 Pt100 sensor DIN class A in 4-wire-circuit


Control technology
	Language setting
	German, English, Spanish, French, Polish, Czech, Hungarian

	ControlCOCKPIT
	SingleDISPLAY. Adaptive multifunctional digital PID-microprocessor controller with high-definition TFT-colour display

	Timer
	Digital backwards counter with target time setting, adjustable from 1 minute to 99 days

	Function SetpointWAIT
	the process time does not start until the set temperature is reached

	Calibration
	three freely selectable temperature values

	adjustable parameters
	temperature (Celsius or Fahrenheit), air flap position, programme time, time zones, summertime/wintertime


Ventilation
	Convection
	natural convection

	Fresh air
	Admixture of pre-heated fresh air by electronically adjustable air flap

	Vent
	vent connection with restrictor flap


Communication
	Documentation
	programme stored in case of power failure

	Programming
	AtmoCONTROL software for reading out, managing and organising the data logger via Ethernet interface (temporary trial version can be downloaded). USB stick with AtmoCONTROL software available as accessory (on demand).


Safety
	Temperature control
	adjustable electronic overtemperature monitor and mechanical temperature limiter TB, protection class 1 according to DIN 12880 to switch off the heating approx. 20°C above nominal temperature

	Autodiagnostic system
	for fault analysis


Standard equipment
	Works calibration certificate
	incl. works calibration certificate for +37°C

	Door
	fully insulated stainless steel door with 2-point locking (compression door lock)

	Door
	inner glass door

	Internals
	2 stainless steel grid(s), electropolished


Stainless steel interior
	Interior
	easy-to-clean interior,made of stainless steel, reinforced by deep drawn ribbing with integrated and protected large-area heating on four sides

	Volume
	161 l

	Dimensions
	w(A) x h(B) x d(C): 560 x 720 x 400 mm

	Max. number of internals
	8

	Max. loading of chamber
	210 kg

	Max. loading per internal
	20 kg


Textured stainless steel casing
	Dimensions
	w(D) x h(E) x d(F): 745 x 1104 x 584 mm (d +56mm door handle)

	Housing
	rear zinc-plated steel


Electrical data
	Voltage Electrical load
	230 V, 50/60 Hz
approx. 1600 W

	Voltage Electrical load
	115 V, 50/60 Hz
approx. 900 W


Ambient conditions
	Set Up
	The distance between the wall and the rear of the appliance must be at least 15 cm. The clearance from the ceiling must not be less than 20 cm and the side clearance from walls or nearby appliances must not be less than 5 cm.

	Altitude of installation
	max. 2,000 m above sea level

	Ambient temperature
	+5 °C to +40 °C

	Humidity rh
	max. 80 %, non-condensing

	Overvoltage category
	II

	Pollution degree
	2


Packing/shipping data
	Transport information
	The appliances must be transported upright

	Customs tariff number
	8419 8998

	Country of origin
	Federal Republic of Germany

	WEEE-Reg.-No.
	DE 66812464

	Dimensions approx incl. carton
	w x h x d: 830 x 1300 x 800 mm

	Net weight
	approx. 96 kg

	Gross weight carton
	approx. 122 kg



