Agricultural Research Council

Plant Health and Protection

Private Bag X134, Queenswood, Pretoria 0121

Reference #: 01_05/26

# REQUEST FOR QUOTATION (RFQ)

## Autoclave Inspection, Service and Pressure Testing

### Client

Agricultural Research Council – Plant Health and Protection
KwaMhlanga/Moloto Road (R573)
Roodeplaat
Pretoria, 0039

### Contact Person

Christa Smit
Mobile: 082 415 2862
Email: Smitc@arc.agric.za

# Scope of Work

The Agricultural Research Council (ARC) hereby invites suitably qualified and accredited service providers to submit quotations for the inspection, servicing, and pressure testing of autoclaves/pressure vessels located at the above-mentioned facility.

### Quantity

- 17 x Pressure Vessels / Autoclaves

# Service Requirements

The following checklist indicates the minimum inspection and service requirements. Bidders are requested to indicate, by means of a tick (✓), which checks and services are included in their quotation. Service providers may also include additional recommendations or inspections deemed necessary.

| Service Item | Included (✓) |
| --- | --- |
| Power switch – verify proper functioning |  |
| Control unit (PLC, battery) – verify proper functioning and condition (if applicable) |  |
| Inspect inner chamber and outer jacket for corrosion |  |
| Check pressure gauges for accuracy |  |
| Check pressure sensors for accuracy |  |
| Inspect piping for leaks |  |
| Clean strainers and refit |  |
| Clean door/lid gasket, apply silicone spray, and refit |  |
| Test and verify safety valve operation |  |
| Verify door/lid locking assembly and inspect locking pins |  |
| Inspect valves (steam, air, water, exhaust), including solenoid valves and valve seats |  |
| Inspect steam generator/boiler, including low water cut-off devices, probes, debris removal, heater tightening, contactors, and descaling requirements |  |
| Check heating elements/coils in boiler |  |
| Inspect electrical connections and protection from steam/moisture |  |
| Inspect contactors/relays for pitting, heat damage, and arcing |  |
| Check steam/thermostatic traps and replace seats/capsules if required |  |
| Inspect vacuum pump, including bearings and worn components |  |
| Inspect door-check springs and retaining bolts |  |
| Other (please specify) |  |

### Comments / Recommendations

# Pressure Testing Requirements

Service providers must indicate compliance with the following requirements and provide supporting accreditation details where applicable.

| Pressure Test Requirement | Included / Details |
| --- | --- |
| SANAS accredited |  |
| SANAS Accreditation Number |  |
| Approved Inspection Authority (AIA) – LUVP 070 approved by Department of Labour (PER 066) and Department of Mineral Resources and Energy (DME 038) |  |
| Approved for manufacturing and independent in-service inspections (PER 0035 INSP) of Pressure Equipment |  |
| Compliance with Occupational Health and Safety Act, 1993 (Act No. 85 of 1993) |  |
| Competent In-Service Inspector (SAQCC – CP-PV & SG) registered with SAQCC and listed on SANAS schedule |  |
| External inspection |  |
| Internal inspection (where possible) |  |
| Wall thickness report |  |
| Pressure testing to required test pressure |  |
| Detailed inspection report and findings |  |
| Pressure Test Certificate issued |  |
| Other accreditation fields (please specify) |  |
| Other (please specify) |  |

### Comments

# Submission Requirements

Quotations should include:

- Company profile and relevant experience

- Accreditation certificates and registration details

- Estimated duration of work

- Pricing breakdown

- Availability and lead times

- Any exclusions or assumptions

# Closing Notes

All work must be carried out in accordance with applicable South African regulations, safety standards, and pressure equipment inspection requirements.

Service providers are encouraged to include any additional inspections, preventative maintenance recommendations, or compliance requirements not specifically listed above.